# Table 1: Information for ETX-260804-0335 under investigation

| Processing Analyst | Reading Analyst | Sample ID | Related Microbial ID | Media with microbial growth | Microbial ID |
| --- | --- | --- | --- | --- | --- |
| Gabrielle Surber | Elysse Nioupin | ETX-260804-0335 | ETX-260904-0331 | 1 x 100mL TSB bottle | Bacillus megaterium |

# Table 2: Environmental Monitoring from Processing Performed on 25Aug26

| Environmental Monitoring (EM) Sampling Site | Frequency |  | Date (DDMMMYY) | Analyst (Initials) |  | Day /Week(s) | Observation* |  | Environmental Monitoring Plate ETX ID | Microbial ID | Notes |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM Bracketing |  |  |  |  |  |  |  |  |  |  |  |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 21Aug26 | GS |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 25Aug26 | GS |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 26Aug26 | GS |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Biological Safety Cabinet EM Bracketing Biological Safety Cabinet (BSC) |  |  |  |  |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 24Aug26 | ES |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 25Aug26 | GS |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 26Aug26 | GS |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 24Aug26 | ES |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 25Aug26 | GS |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 26Aug26 | GS |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Weekly Active Air Sampling Bracketing |  |  |  |  |  |  |  |  |  |  |  |  |
| Active Air Sampling of Cleanrooms |  | Weekly | 25Aug26 | SMO | Week on Testing Date |  |  | 4 CFU (ISO 8 114) | ETX-260901-0112 | 3 Gram (+) cocci, 1 Hyphae |  | None |
| Surface Sampling of Anteroom and Cleanroom Bracketing |  |  |  |  |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms |  | Weekly | 25Aug26 | SMO | Week on Testing Date |  |  | No growth | N/A | N/A |  | None |