Phase I

| # Section A – Test Details |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| # Initiator: | # Sonal Uprety # (Written by: Qiyue Chen) | # Test Date: |  | # 07Aug26 |  |  |  | # Date Initiated: |  | # 07Aug26 |  |
| # Date of Incident: | # 07Aug26 | # Test Name: |  | # Scan RDI Sterility Test |  |  |  | # Sample ID: |  | # ETX-260805-0189 # GoGoMeds Select (E10747) |  |
| # Name of Analyst who Performed the Test: | # Prepper: # Elysse Nioupin (EN) # Processor: Sonal Uprety (SU) # Changeover Processor: # Sonal Uprety (SU) # Reader: # Varsha Subramanian (VV) | # Sample / Active Name |  | # Tirzepatide/B12 25/1mg/mL Injectable |  |  |  | # Lot #: |  | # 18468 |  |
|  |  | # Dosage Form |  | # Injectable |  |  |  |  |  |  |  |
| # Description of # Incident: | # On 07Aug26, sample ETX-260805-0189 was found positive for viable microorganisms after ScanRDI testing. |  |  |  | # Client Care notified of OOS # Sterility and Endotoxin test related OOS; Client must be notified within 24 hours of OOS |  |  |  |  |  |  |
|  | # SOP / Test Method #: # 2.600.023 (10) # 2.700.004 (03) | # Effective Date: # 07Apr25 # 02Jan23 |  |  |  |  | # SOP / Test Method Rev: 10 # Rev: 03 |  | # Limits/Specification: # <1 event |  |  |
| # Manager notified | # Manager/designee name: # Kathan Parikh | # Date: # 07Aug26 |  |  |  |  | # Quality Assurance Department Notified: |  | # Yes No |  |  |
| # Section B – Investigation and Root Cause Analysis |  |  |  |  |  |  |  |  |  |  |  |
| # Personnel Discrepancy |  |  |  |  |  | # Comments |  |  |  |  |  |
| # Analyst interviewed? |  | # Yes No N/A |  |  |  | # Yes, analysts Elysse Nioupin, Sonal Uprety and Varsha Subramanian were interviewed comprehensively. |  |  |  |  |  |
| # Correct samples analyzed? |  | # Yes No N/A |  |  |  | # Yes, ETX-260805-0189 |  |  |  |  |  |
| # Correct SOP / Test Method used/followed? |  | # Yes No N/A |  |  |  | # Yes, as per SOP 2.600.023, 2.700.004 |  |  |  |  |  |
| # Correct sample collection/technique? |  | # Yes No N/A |  |  |  | # Yes, as per SOP 2.600.023, 2.700.004 |  |  |  |  |  |
| # Are the Test Result Records Attached & Calculations Verified and Correct? |  | # Yes No N/A |  |  |  | # Yes, See 080726-1040-2 for more information. |  |  |  |  |  |
| # Is the analyst qualified to perform the test? |  | # Yes No N/A |  |  |  | # Yes, all analysts are trained and qualified by quality to perform the test. |  |  |  |  |  |
| # Is formula worksheet provided and verified? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  |  |
| # Does the processing method/ method set conform to the required specification? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  |  |
| # Was the sample stored appropriately? |  | # Yes No N/A |  |  |  | # Yes, Information is available in Eagle Trax Sample Location History under ETX-260805-0189 |  |  |  |  |  |
| # Material Discrepancy |  |  |  |  |  | # Lot Number |  |  |  |  | # Expiry Date |
| # Is the integrity of the sample(s) questionable? |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |
| # Are the other test results from this run under investigation? |  | # Yes No N/A (See Phase I summary) |  |  |  |  |  |  |  |  |  |
| # Appropriate /correct glassware used? |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |
| # Are reagents stored in proper condition and not expired? |  | # Yes No N/A |  |  |  | # Scan Consumables: # See the attached data packet # Environmental Plates: # TSA and Surface Plate: see attached environmental logs # Monthly Cleaning: # H2O2 strips and IPA: See attached monthly cleaning logs |  |  |  |  | # Scan Consumables: # See attached data packet # Environmental Plates: # TSA and Surface Plate: see attached environmental logs # Monthly Cleaning: # H2O2 strips and IPA: See attached monthly cleaning logs |
| # Standard Used |  | # A. brasiliensis |  |  |  | # 05222026-16404-AB |  |  |  |  | # 22May28 |
| # Solvent Used |  | # Not Applicable |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Is the standard the source of the root cause? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Has the analyst been instructed to keep the original test solution for possible retesting? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Equipment Discrepancy |  |  |  |  |  | # Equipment ID |  |  |  |  | # Calibration Due Date |
| # Was the instrument set up correctly? |  | # Yes No N/A |  |  |  | # E001040 |  |  |  |  | # Oct26 |
| # Is the clean room facility certified? |  | Yes No N/A |  |  |  | # E001737 (CR115) |  |  |  |  | # Dec25 |
| # System Performance Check (if applicable)? |  | # Yes No N/A |  |  |  | # E001040 |  |  |  |  | # Oct26 |
| # System Suitability RSD < 2% (PDA)? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # System Suitability RSD < 10% (QDA,CAD)? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Do the environmental monitoring results conform to required standards? |  | Yes No N/A |  |  |  | # See Phase I Summary |  |  |  |  | # See Phase I Summary |
| # Do the controls (i.e., positive, negative, etc.) conform to the required standards? |  | Yes No N/A |  |  |  | # See Phase I Summary |  |  |  |  | # See Phase I Summary |
| # Is equipment/instrument functional and calibrated/qualified/certified/PM? |  | # Yes No N/A |  |  |  | # Incubator E001034 # (Sensor E001501) # Incubator E001031 (Sensor E001505) |  |  |  |  | # Aug26 # Feb26 # Aug26 # Feb26 |
| # Are chromatogram/spectra/raw data correct? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Section C – Phase I Summary |  |  |  |  |  |  |  |  |  |  |  |
| # Has root cause been identified as a possible laboratory error? # Note: Root cause must be conclusive, verifiable, and well-documented. # Yes # If retesting is required, complete Phase II of the OOS (initiate Form 3.100.019.F02). # No, the OOS result is valid. |  |  |  |  |  |  |  |  |  |  |  |
| # Phase I Summary: # All analysts involved in the prepping, processing, and reading of the samples – Elysse Nioupin, Sonal Uprety and Varsha Subramanian – were interviewed and their answers are recorded throughout this document. # The sample was stored upon arrival according to the Client’s instructions. Analysts Elysse Nioupin and Sonal Uprety confirmed the integrity of the samples throughout both the preparation and processing stages. No leaks or turbidity were observed at any point, verifying the integrity of the sample. # All reagents and supplies mentioned in the material section above were stored according to the suppliers’ recommendations, and their integrity was visually verified before utilization. Moreover, each reagent and supply had valid expiration dates. # During the preparation phase, Elysse Nioupin disinfected the samples using acidified bleach and placed them into a pre-disinfected storage bin. On 07Aug26, prior to sample processing, Sonal Uprety performed a second disinfection with acidified bleach, allowing a minimum contact time of 10 minutes before transferring the samples into the cleanroom suites. A final disinfection step was completed immediately before the samples were introduced into the ISO 5 Biological Safety Cabinet (BSC), E001313, located within the middle ISO 7 buffer room, (Suite 115A), All activities were performed in accordance with SOP 2.600.023, Rapid Scan RDI® Test Using FIFU Method. # The cleanroom used for testing and changeover procedures (Suite 115) comprises three interconnected sections: the innermost ISO 7 cleanroom (115B), which connects to the middle ISO 7 buffer room (115A), and then to the outermost ISO 8 anteroom (115). A positive air pressure system is maintained throughout the suite to ensure controlled, unidirectional airflow from 115B through 115A and into 115. The ISO 5 BSC E001313, located in the middle ISO 7 buffer room, (Suite 115A), was used for both testing and changeover steps. It was thoroughly cleaned and disinfected prior to each procedure in accordance with SOP 2.600.018 (Cleaning and Disinfecting Procedure for Microbiology). Additionally, BSC E001313 was certified and approved by both the Engineering and Quality Assurance teams. Sample processing and changeover were conducted in the ISO 5 BSC E001313 in the middle ISO 7 buffer room, (Suite 115A) by Sonal Uprety on 07Aug26. # The analyst, Varsha Subramanian, confirmed that the equipment was set up as per SOP 2.700.004 (Scan RDI® System – Operations (Standard C3 Quality Check and Microscope Setup and Maintenance), and the negative control and the positive control for the analyst, Varsha Subramanian, yielded expected results. # On 07Aug26, a rapid sterility test was conducted on the sample using the ScanRDI method. The sample was initially prepared by Analyst Elysse Nioupin, processed by Sonal Uprety, and subsequently read by Sonal Uprety. The test revealed 4 Rod-shaped viable microorganisms, see table 1. # Table 2 (see attached table) presents the environmental monitoring results for ETX-260805-0189. The environmental monitoring (EM) plates were incubated for no less than 48 hours at 30-35°C and no less than an additional five days at 20-25°C as per SOP 2.600.002 (Environmental Monitoring of the Clean-room Facility). # Upon analyzing the environmental monitoring results, no microbial growth was observed in surface sampling. However, microbial growth was observed during personnel sampling (right touch) and settling plates on the date of testing, and weekly active air sampling and weekly surface sampling from the week of testing. Specifically, 1 CFU on RT was detected during personnel sampling and was submitted for microbial identification under sample ID ETX-260817-0447, where the organism was identified as Micrococcus luteus (plate was desiccated on 5 day read). Additionally, 2 CFUs on sett 2 were detected during settling plates and was submitted for differential staining under sample ID ETX-260817-0507, where the organisms were identified as Insufficient read Gram (+) rods (plate was desiccated on 5 day read). Furthermore, 1 CFU in ISO 8 115 was detected during weekly active air sampling under sample ID ETX-260817-0370 (Gram (+) short rods). Also, 2 CFUs on floor in 115A ISO 7 were detected during weekly surface sampling under sample ID ETX-260817-0366 (Bacillus megaterium). # Monthly cleaning and disinfection, using H₂O₂, of the cleanroom (ISO 7) and its containing Biosafety Cabinets (BSCs, ISO 5) were performed on 26Jul26, as per SOP 2.600.018 Cleaning and Disinfection Procedure. It was documented that all H₂O₂ indicators passed. # Analyzing a 6-month sample history for GoGoMeds Select (E10747), this specific analyte "Tirzepatide/B12 25/1mg/mL Injectable" has had no prior failures using the Scan RDI method during this period. # To assess the potential for sample-to-sample contamination contributing to the positive results, a comprehensive review was conducted of all samples processed on the same day. All other samples processed by the analyst and other analysts that day tested negative. These findings suggest that cross-contamination between samples is highly unlikely. # Based on the observations outlined above, it is unlikely that the failing results were due to reagents, supplies, the cleanroom environment, the process, or analyst involvement. Consequently, the possibility of laboratory error contributing to this failure is minimal and the original result is deemed to be valid. |  |  |  |  |  |  |  |  |  |  |  |
| # Reviewer Comments: |  |  |  |  |  |  |  |  |  |  |  |
| # Acknowledgement and Closure |  |  |  |  |  |  |  |  |  |  |  |
| # Can investigation be closed? # Yes # Original result is valid. # No, initiate Form 3.100.019.F02. |  |  |  |  |  |  |  |  |  |  |  |
| # Prepared by (Print): |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |
| # Lab Manager/designee (Print): |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |
| # Approved by QA Representative (Print): # Closed: YES NO |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |

# Table 1: Information for ETX-260805-0189 under investigation

| Processing Analyst | Reading Analyst | Sample ID | Events | Confirmed Microbial Events | Morphology Description |
| --- | --- | --- | --- | --- | --- |
| Sonal Uprety | Varsha Subramanian | ETX-260805-0189 | 187 | 4 | Rod-shaped Morphology |

# Table 2: Environmental Monitoring from Testing Performed on 07Aug26

| Environmental Monitoring (EM) Sampling Site |  | Frequency | Date (DDMMM YYYY) | Analyst (Initials) |  | Day /Week(s) |  | Observation |  | Environmental Monitoring Plate ETX ID |  | Microbial ID | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM Bracketing (Sonal Uprety) |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Personal (Left Touch and Right Touch) |  | Daily | 07Aug26 | SU |  | Date of Testing |  | 1 CFU on RT |  | ETX-260817-0447 |  | Micrococcus luteus | None |
| Biological Safety Cabinet EM Bracketing Biological Safety Cabinet (BSC) E001313 and E001313 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 E001313 (4 locations) |  | Daily | 07Aug26 | SU |  | Date of Testing |  | No growth |  | N/A |  | N/A | None |
| Surface Sampling of ISO 5 E001313 (4 locations) |  | Daily | 07Aug26 | SU |  | Date of Testing |  | No growth |  | N/A |  | N/A | None |
| Settling Sampling of ISO 5 E001313 (2 locations) |  | Daily | 07Aug26 | SU |  | Date of Testing |  | 2 CFUs on sett 2 |  | ETX-260817-0507 |  | Insufficient read Gram (+) rods | None |
| Settling Sampling of ISO 5 E001313 (2 locations) |  | Daily | 07Aug26 | SU |  | Date of Testing |  | 2 CFUs on sett 2 |  | ETX-260817-0507 |  | Insufficient read Gram (+) rods | None |
| Weekly Active Air Sampling Bracketing – Cleanroom 115 – CR1737 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Environmental Monitoring (EM) Sampling Site | Frequency |  | Date (DDMMM YYYY) |  | Analyst (Initials) |  | Observation |  | Environmental Monitoring Plate ETX ID |  | Microbial ID |  | Notes |
| Active Air Sampling of Cleanrooms | Weekly |  | 06AUG 2026 |  | SMO |  | 1 CFU in ISO 8 115 |  | ETX-260817-0370 |  | Gram (+) short rods |  | None |
| Surface Sampling of Anteroom and Cleanroom Bracketing - Cleanroom 115 – CR1737 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms | Weekly |  | 06AUG 2026 |  | SMO |  | 2 CFUs on floor in 115A ISO 7 |  | ETX-260817-0366 |  | Bacillus megaterium |  | None |