# Table 1: Information for ETX-260805-0189 under investigation

| Processing Analyst | Reading Analyst | Sample ID | Events | Confirmed Microbial Events | Morphology Description |
| --- | --- | --- | --- | --- | --- |
| Sonal Uprety | Varsha Subramanian | ETX-260805-0189 | 187 | 4 | Rod-shaped Morphology |

# Table 2: Environmental Monitoring from Testing Performed on 07Aug26

| Environmental Monitoring (EM) Sampling Site | Frequency | Date (DDMMM YYYY) | Analyst (Initials) | Day /Week(s) | Observation* | Environmental Monitoring Plate ETX ID | Microbial ID | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM Bracketing |  |  |  |  |  |  |  |  |
| Personal (Left Touch and Right Touch) | Daily | 07AUG 2026 | SU | Date of Testing Scan | 1 CFU on right touch | ETX-260817-0447 | Micrococcus luteus | None |
| Personal (Left Touch and Right Touch) | Daily | 07AUG 2026 | GA | Date of Testing ScanC/O | No growth | N/A | N/A | None |
| Biological Safety Cabinet EM Bracketing Biological Safety Cabinet (BSC) 1313 and 1937 |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 (4 locations) | Daily | 07AUG 2026 | SU | Date of TestingScan | No growth | N/A | N/A | None |
| Surface Sampling of ISO 5 (4 locations) | Daily | 07AUG 2026 | GA | Date of Testing ScanC/O | No growth | N/A | N/A | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 07AUG 2026 | SU | Date of TestingScan | 2 CFUs on Sett1 | ETX-260817-0507 |  | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 07AUG 2026 | GA | Date of Testing ScanC/O | No growth | N/A | N/A | None |
| Weekly Active Air Sampling Bracketing 115 |  |  |  |  |  |  |  |  |
| Active Air Sampling of Cleanrooms | Weekly | 06AUG 2026 | SMO | Week of testing Scan | 1 CFU on ISO 8 115 | ETX-260817-0370 | Gram (+) short rods | None |
| Weekly Active Air Sampling Bracketing L-suite |  |  |  |  |  |  |  |  |
| Active Air Sampling of Cleanrooms | Weekly | 07AUG 2026 | SMO | Week of testing ScanC/O | 6 CFUs on 143 SecI air ISO 8 32 CFUs on 143 SecII air ISO 8, and 12 CFUs on 142 air ISO 8 | ETX-260817-0416 ETX-260817-0422 ETX-260817-0429 | Staphylococcus epidermidis Staphylococcus capitis Micrococcus luteus Kocuria palustris Corynebacterium segmentosum Staphylococcus hominis Staphylococcus aureus Corynebacterium aurimucosum Kytococcus sedentarius | None |
| Surface Sampling of Anteroom and Cleanroom Bracketing 115 |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms | Weekly | 06AUG 2026 | SMO | Week of testing Scan | 2 CFU on floor in ISO 7 115A | ETX-260817-0366 | Bacillus megaterium | None |
| Weekly Active Air Sampling Bracketing L-suite |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms | Weekly | 07AUG 2026 | SMO | Week of testing ScanC/O | No growth | N/A | N/A | None |