Phase II

| # Section A – Sample Detail |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| # Sample Name: | # GLP3R/Cagrilinitide |  | # Re-Test Date: |  |  | # 15-Sep-2026 |  |  | # Original Result: |  |  |  | # Fail |  |
| # Name of Analyst who Performed Re-Test: | # Prepping Analyst: Andrew Carrillo (ANC) Processing Analyst: Ayomide Odugbesi (AOD) Reading Analyst: Andrew Carrillo (ANC) |  | # Sample ID: |  |  | # Original test: ETX-260902-0505 Retest: ETX-260914-0470 |  |  | # Re-Test Result: |  |  |  | # Fail |  |
| # Section B – Investigation and Root Cause |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| # Investigation | # Is the test conducted in accordance with the same method /process? |  |  | # Yes No |  |  |  | # SOP#: |  |  |  | # If the answer is NO to any of the investigation item(s), mark the item(s) with sequential number starting with one (1) and comment here: |  |  |
|  | # Is reagents (reference standards, solvent, etc.) stored in proper condition and not expired? |  |  | # Yes No |  |  |  | Lot #: |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  | # Exp. Date: |  |  |  |  |  |  |
|  | # Equipment/Instruments Used: |  |  | # Are they functional and calibrated/PM? |  |  |  | # Calibration / PM due date |  |  |  |  |  |  |
|  | # ISO 5 BSC E001316, ISO 5 BSC E001314 |  |  | # Yes No |  |  |  |  |  |  |  |  |  |  |
|  | # CR 114 (E001736) and CR 115 (E001737) |  |  | # Yes No |  |  |  |  |  |  |  |  |  |  |
|  | # Incubator E001356, E001357, E001034, E001031 |  |  | # Yes No |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  | # Yes No |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  | # Yes No |  |  |  |  |  |  |  |  |  |  |
|  | # Is the clean room facility certified? |  |  | # Yes No N/A |  |  |  | # Certified due date: |  |  |  |  |  |  |
|  | # Do the controls (i.e. positive, negative, C3, etc.,) conform to the required standards? |  |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |  |
|  | # Do the environmental monitoring results conform to required standards? |  |  | # Yes No N/A |  |  |  | # SOP#: |  |  |  |  |  |  |
|  | # System Performance Check (if applicable) |  |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |  |
|  | # System Suitability RSD < 2% (PDA) |  |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |  |
|  | # System Suitability RSD <10 % (QDA,CAD) |  |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |  |
|  | # Is the analyst trained to perform the test? |  |  | # Yes No |  |  |  |  |  |  |  |  |  |  |
|  | # Are the Test Result Records Attached & Calculations Verified and Correct? |  |  | # Yes No |  |  |  |  |  |  |  |  |  |  |
|  | # Does the processing method/ method set conform to the required specification? |  |  | # Yes No |  |  |  |  |  |  |  |  |  |  |
|  | # Re-Testing from the same vial? |  |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |  |
|  | # Is formula worksheet provided and verified? |  |  | # Yes No |  |  |  |  |  |  |  |  |  |  |
| # Section C – Evaluation of Laboratory Testing/Hypothesis Testing (if applicable) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| # N/A |  |  |  |  | # Re-vialing of original working sample solution |  |  |  |  |  |  |  |  |  |
| # Re-injection of sample vial |  |  |  |  | # Re-dilution of a working sample solution from the original stock |  |  |  |  |  |  |  |  |  |
| # Re-injection of original standard vial |  |  |  |  | # Re-dilution of a working standard solution from the original stock standard |  |  |  |  |  |  |  |  |  |
| # Re-vialing of original working standard solution |  |  |  |  | # Other: _________________________________ |  |  |  |  |  |  |  |  |  |
| # Section D – Phase II Summary |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| # All the analysts involved in the testing - the initial prepping analyst (Ishita Sharma), the initial processing analyst (Alex Saravia), the initial reading analyst (Elysse Nioupin), the retest prepping and reading analyst (Andrew Carrillo), and the retest processing analyst (Ayomide Odugbesi) - were interviewed comprehensively. Their answers are recorded throughout this document. This investigation was performed in accordance with MICRO-SOP-53, Sterility Test Out-of-Specification Investigation Procedure. INITIAL TEST UNDER SUBMISSION ETX-260902-0505: All analysts involved in the prepping, processing, and reading of the initial sample – Ishita Sharma, Alex Saravia, and Elysse Nioupin – were interviewed comprehensively. Upon arrival, the sample was stored refrigerated in accordance with the Client’s instructions. Analysts verified the integrity of the sample throughout both preparation and processing stages. No leaks or turbidity were observed at any point, verifying sample integrity. All reagents and supplies mentioned in the material section above were stored according to the suppliers’ recommendations, and their integrity was visually verified before utilization. Moreover, all reagents and supplies had valid expiration dates. The functionality of all equipment was confirmed by reviewing data generated by our comprehensive in-house continuous monitoring system. The cleanroom suite (CR 114) used for processing procedures comprises three interconnected sections: the innermost ISO 7 cleanroom (114B), which connects to the middle ISO 7 buffer room (114A), and then to the outermost ISO 8 anteroom (114). A positive air pressure system is maintained throughout the suite to ensure controlled, unidirectional airflow from 114B through 114A and into 114. The ISO 5 Biological Safety Cabinet (BSC) E001316, located in the innermost ISO 7 room, (114B), was used for initial sample processing steps. It was thoroughly cleaned and disinfected prior to each procedure in accordance with MICRO SOP-9 (Cleaning and Disinfecting Procedure for Microbiology). Additionally, BSC E001316 was certified and approved by both the Engineering and Quality Assurance teams. Initial sample processing was conducted in the ISO 5 BSC E001316 in the innermost ISO 7 room, (114B) by Alex Saravia on 04Sep26 as per MICRO-SOP-5 (USP 71> / EP 2.6.1 Sterility Test). On 04Sep26, the sample vials for ETX-260902-0505 were received from Sample Submissions and brought into the Sterile Microbiology lab following multi-stage disinfection with acidified bleach (minimum 10-minute contact time at each transfer boundary). In ISO 5 BSC E001316, the containers were aseptically opened, and membrane filtration was performed (45 gm filtered per media across 3 vials) as per MICRO-SOP-5. Media bottles were transferred into designated incubators E001356 and E001357. On 10Sep26 (Day 6 of incubation), microbial growth was observed in 1 x 100mL TSB media bottle by analyst Elysse Nioupin and confirmed by supervisor Robin Seymour. The positive TSB bottle was submitted for Differential Staining and Microbial Identification under ETX-260910-0290, which definitively identified the isolate as Microbacterium sp. PM5 (Gram-positive rods). All concurrent negative controls for the membrane filtration method remained sterile (no growth). Upon analyzing environmental monitoring results, no microbial growth was detected on operator personnel fingertip touch plates, settling plates, or ISO 5 BSC surface contact plates on the date of testing (04Sep26), the date before (03Sep26), or the subsequent testing date (08Sep26). Weekly active air monitoring of ISO 8 Room 114 recovered 2 CFUs (ETX-260914-0487, identified as Corynebacterium ureicelerivorans and Mycobacterium grossiae), which differed in morphology and genus from the sample isolate and was physically isolated within the outermost ISO 8 anteroom. Monthly cleaning and disinfection of Suite CR114 was performed on 30Aug26 per MICRO-SOP-9 with all H2O2 indicators passing. All other samples processed in the same testing session tested negative, ruling out cross-contamination. However, when investigating the lot history during the Phase I investigation, the laboratory discovered that the same lot (Lot 2608-216) had previously been submitted for ScanRDI sterility testing under ETX-260807-0602 (Test Record 081326-2017-2) and yielded a passing result (Count: 14/0, Pass). Based on this conflicting passing ScanRDI result for the same lot, the laboratory requested additional sample vials from the client for retesting under ETX-260914-0470 and initiated a Phase II investigation (Form 3.100.019.F02). # Phase II Summary (continued): # PHASE II RETEST UNDER SUBMISSION ETX-260914-0470: In accordance with Phase II investigation procedures under Form 3.100.019.F02, following the laboratory's discovery of the prior passing ScanRDI result (ETX-260807-0602) during the sample history investigation, additional sample vials from lot 2608-216 were requested from the client and submitted for retesting under submission ETX-260914-0470. The retest sample was prepared by Andrew Carrillo on 14Sep26 and processed by Ayomide Odugbesi on 15Sep26 in accordance with MICRO-SOP-5 (USP 71> / EP 2.6.1 Sterility Test) via membrane filtration (60 gm filtered across 4 vials). The retest was conducted in certified ISO 5 BSC E001314 within Cleanroom Suite 115 following standard multi-barrier disinfection and strict aseptic technique. All culture media (FTM Lot 2567320, TSB Lot 2567280) and reagents (Fluid D Lot 687774, Bacteriostatic Water Lot NM7681) were within valid expiration dates. Media bottles were incubated in designated incubators E001356 and E001357. Negative controls were handled concurrently and remained sterile. On 21Sep26 (Day 06 of incubation), microbial growth was observed in 1 x 100mL TSB media bottle by analyst Andrew Carrillo and confirmed by supervisor Robin Seymour. The positive bottle was submitted for Microbial Identification under ETX-260921-0498. Confirmatory identification definitively identified the isolate as Microbacterium sp. PM5 (Gram-positive short rods), matching the identical species and strain designation (PM5), morphology (white shiny lawn growth), media selection (TSB), and incubation kinetics (Day 6) of the initial failure under ETX-260902-0505 (ETX-260910-0290). All other samples processed in the batch on 15Sep26 tested negative. Upon analyzing environmental monitoring results for the retest processing on 15Sep26, no microbial growth was detected on operator personnel fingertip touch plates, settling plates, or ISO 5 BSC surface contact plates on the date of testing (15Sep26), the preceding testing date (14Sep26), or the subsequent testing date (16Sep26). Weekly active surface monitoring of Cleanroom Suite 115 showed no growth; weekly active air sampling in the outermost ISO 8 cleanroom 115 recovered 9 CFUs (ETX-260923-0402), identified as Kocuria indica, Brevibacterium sp. CS2, Corynebacterium sp, Micrococcus luteus, Paracoccus yeei, Staphylococcus hominis, and Kocuria rhizophila. 100% of these environmental isolates differed distinctly in genus and morphology from the product contamination isolate (Microbacterium sp. PM5). The ISO 5 critical zones in BSC E001314 remained completely sterile (0 CFU) and in optimal control throughout testing, with positive pressure cascades preventing anteroom air ingress. (Refer to Table 1 for sample details, Table 2 for environmental monitoring from processing performed on 04Sep26, and Table 3 for environmental monitoring from processing performed on 15Sep26). EVALUATION OF RETEST FINDINGS JUSTIFICATION OF PRIOR SCANRDI RESULT: The replication of microbial recovery in the independent retest—conducted by a different processing analyst (Ayomide Odugbesi) in a different cleanroom suite (Cleanroom Suite 115 vs. Suite 114) and a different prepping/reading analyst (Andrew Carrillo)—conclusively confirms that the microbial recovery was not an isolated laboratory anomaly, as the identical microorganism (Microbacterium sp. PM5) was recovered on the identical incubation day (Day 6) in the same media type (TSB) in both tests. Importantly, the investigation evaluated why the prior rapid ScanRDI sterility test conducted on this lot (submission ETX-260807-0602, test record 081326-2017-2) yielded a passing result (Count 14/0, Pass). First, per EagleTrax records for ETX-260807-0602, only 1 single container (Number of Items in Batch: 1, 30 mL/gm) was submitted and tested via ScanRDI, whereas USP 71> compendial testing analyzed 3 vials (45 gm filtered per media) for initial test ETX-260902-0505 and 4 vials for retest ETX-260914-0470. Second, official records confirm that 'No Method Suitability on file' exists for ScanRDI with this specific peptide analyte formulation (Cagrilinitide / Retatrutide). Third, in compounded parenteral formulations, microbial contamination at low bioburden levels is well recognized to exhibit a non-uniform, heterogeneous distribution across individual batch containers. Under single-vial sampling (n = 1), random chance naturally selects a container devoid of viable cells—as occurred with the single vial tested under ETX-260807-0602—whereas multi-container sampling captures vials harboring viable cells. When viable cells were introduced into 14-day compendial broth enrichment (USP 71> membrane filtration), the organisms were successfully enriched and recovered in TSB on Day 6 in both independent testing events (ETX-260902-0505 and ETX-260914-0470), yielding the identical strain (Microbacterium sp. PM5). Therefore, the passing ScanRDI result and the two failing USP 71> compendial results are scientifically consistent and attributable to non-uniform distribution of low-concentration microbial bioburden within the compounded lot. Consequently, inherent product contamination in Lot 2608-216 is confirmed, the initial failing result for ETX-260902-0505 is valid, retest ETX-260914-0470 is confirmed failing, and Lot 2608-216 fails sterility requirements. Inherent product contamination in Lot 2608-216, confirmed by the replication of identical microbial recovery kinetics (Day 6 in TSB) and the identical strain (Microbacterium sp. PM5) in retest ETX-260914-0470. The discrepancy with the passing ScanRDI result (ETX-260807-0602) is scientifically attributed to non-uniform microbial distribution at low concentrations within the compounded batch. Original result confirmed valid. # Phase II Summary (continued): |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| # Has root cause been identified? # Yes No | # The most probable cause is: # Client Care notified: Yes No |  |  |  |  |  |  |  |  |  |  |  |  |  |
| # Section E – Corrective and Preventive Action |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| # Corrective Action(s): |  |  |  |  |  |  |  |  |  |  | # OOS cause (select all) |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  | # Equipment Error |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  | # Procedure |  |  | # Analyst |
|  |  |  |  |  |  |  |  |  |  |  | # Training |  |  | # System |
|  |  |  |  |  |  |  |  |  |  |  | # External Phenomena |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  | # Other: ____________ |  |  |  |
| # Preventive Action(s) |  |  |  |  |  |  |  |  |  |  | # CAPA required: # Yes, CAPA # _________ # No |  |  |  |
| # FOLLOW-UP REQUIRED: Yes No |  |  |  |  |  |  | # FOLLLOW-UP DATE: |  |  |  |  |  |  |  |
| # Section F – Acknowledgement and Closure |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| # Prepared by: |  | # Print Name: |  |  |  |  |  |  |  | # Date: |  |  |  |  |
| # Reviewed by: |  | # Print Name: |  |  |  |  |  |  |  | # Date: |  |  |  |  |
| # Approved by (QA): |  | # Print Name: |  |  |  |  |  |  |  | # Date: |  |  |  |  |

# Table 1: Information for ETX-260902-0505 and ETX-260914-0470 under investigation

| Processing Analyst | Reading Analyst | Sample ID | Related Microbial ID | Media with microbial growth | Microbial ID |
| --- | --- | --- | --- | --- | --- |
| Alex Saravia | Elysse Nioupin | ETX-260902-0505 | ETX-260910-0290 | 1 x 100mL TSB bottle | Microbacterium sp. PM5 (Gram (+) rods) |
| Ayomide Odugbesi | Andrew Carrillo | ETX-260914-0470 | ETX-260921-0498 | 1 x 100mL TSB bottle | Microbacterium sp. PM5 (Gram (+) short rods) |

# Table 2: Environmental Monitoring from Processing Performed on 04Sep26

| Environmental Monitoring (EM) Sampling Site | Frequency |  | Date (DDMMMYY) | Analyst (Initials) |  | Day /Week(s) | Observation* |  | Environmental Monitoring Plate ETX ID | Microbial ID | Notes |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM |  |  |  |  |  |  |  |  |  |  |  |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 03Sep26 | ES |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 04Sep26 | ES |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 08Sep26 | ES |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Biological Safety Cabinet EM for E001316 |  |  |  |  |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 03Sep26 | ES |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 04Sep26 | ES |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 08Sep26 | ES |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 03Sep26 | ES |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 04Sep26 | ES |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 08Sep26 | ES |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Weekly Active Air Sampling of Cleanroom Suite 114 (E001736) with Processing BSC |  |  |  |  |  |  |  |  |  |  |  |  |
| Active Air Sampling of Cleanrooms |  | Weekly | 04Sep26 | ISS | Week of Testing |  |  | 2 CFUs (ISO 8 114) | ETX-260914-0487 | Corynebacterium ureicelerivorans Mycobacterium grossiae |  | None |
| Weekly Surface Sampling of Cleanroom Suite 114 (E001736) with Processing BSC |  |  |  |  |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms |  | Weekly | 04Sep26 | ISS | Week of Testing |  |  | No growth | N/A | N/A |  | None |

Table 3: Environmental Monitoring from Processing Performed on 15Sep26

| Environmental Monitoring (EM) Sampling Site | Frequency |  | Date (DDMMMYY) | Analyst (Initials) |  | Day /Week(s) | Observation* |  | Environmental Monitoring Plate ETX ID | Microbial ID | Notes |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM |  |  |  |  |  |  |  |  |  |  |  |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 14Sep26 | AOD |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 15Sep26 | AOD |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 16Sep26 | AOD |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Biological Safety Cabinet EM for E001314 |  |  |  |  |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 14Sep26 | AOD |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 15Sep26 | AOD |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 16Sep26 | AOD |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 14Sep26 | AOD |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 15Sep26 | AOD |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 16Sep26 | AOD |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Weekly Active Air Sampling of Cleanroom Suite 115 (E001737) with Processing BSC |  |  |  |  |  |  |  |  |  |  |  |  |
| Active Air Sampling of Cleanrooms |  | Weekly | 15Sep26 | SMO | Week of Testing |  |  | 9 CFUs (ISO 8 115) | ETX-260923-0402 | Kocuria indica Brevibacterium sp. CS2 Corynebacterium sp Micrococcus luteus Paracoccus yeei Staphylococcus hominis Kocuria rhizophila |  | None |
| Weekly Surface Sampling of Cleanroom Suite 115 (E001737) with Processing BSC |  |  |  |  |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms |  | Weekly | 15Sep26 | SMO | Week of Testing |  |  | No growth | N/A | N/A |  | None |