Phase I

| # Section A – Test Details |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| # Initiator: | # {{ analyst_name }} # (Written by: Qiyue Chen) | # Test Date: |  | # {{ process_date }} |  |  |  | # Date Initiated: |  | # {{ test_date }} |  |
| # Date of Incident: | # {{ test_date }} | # Test Name: |  | # USP <71> Sterility Test |  |  |  | # Sample ID: |  | # {{ sample_id }} # {{ client_name }} |  |
| # Name of Analyst who Performed the Test: | # Prepper: # {{ prepper_name }} ({{ prepper_initial }}) # Processor: {{ analyst_name }} ({{ analyst_initial }}) # Aliquoting Analyst: # {{ aliquoting_name }} ({{ aliquoting_initial }}) | # Sample / Active Name |  | # {{ sample_name }} |  |  |  | # Lot #: |  | # {{ lot_number }} |  |
|  |  | # Dosage Form |  | # {{ dosage_form}} |  |  |  |  |  |  |  |
| # Description of # Incident: | # On {{ test_date }}, sample {{ sample_id }} was found positive for viable microorganisms after USP <71> sterility testing. |  |  |  | # Client Care notified of OOS # Sterility and Endotoxin test related OOS; Client must be notified within 24 hours of OOS |  |  |  |  |  |  |
|  | # SOP / Test Method #: # 2.600.059 (01) | # Effective Date: # 03Jan25 |  |  |  |  | # SOP / Test: # Rev: 01 |  | # Limits/Specification: # <1 event |  |  |
| # Manager notified | # Manager/designee name: # Kathan Parikh | # Date: # {{ test_date }} |  |  |  |  | # Quality Assurance Department Notified: |  | # Yes No |  |  |
| # Section B – Investigation and Root Cause Analysis |  |  |  |  |  |  |  |  |  |  |  |
| # Personnel Discrepancy |  |  |  |  |  | # Comments |  |  |  |  |  |
| # Analyst interviewed? |  | # Yes No N/A |  |  |  | # Yes, analysts {{ prepper_name }}, {{ analyst_name }}, and {{ aliquoting_name }} were interviewed comprehensively. |  |  |  |  |  |
| # Correct samples analyzed? |  | # Yes No N/A |  |  |  | # Yes, sample ID: {{ sample_id }} |  |  |  |  |  |
| # Correct SOP / Test Method used/followed? |  | # Yes No N/A |  |  |  | # Yes, as per SOP# 2.600.059 |  |  |  |  |  |
| # Correct sample collection/technique? |  | # Yes No N/A |  |  |  | # Yes, as per SOP# 2.600.059 |  |  |  |  |  |
| # Are the Test Result Records Attached & Calculations Verified and Correct? |  | # Yes No N/A |  |  |  | # Yes, Yes, Information is available in EagleTrax under {{ sample_id }} |  |  |  |  |  |
| # Is the analyst qualified to perform the test? |  | # Yes No N/A |  |  |  | # Yes, analysts {{ prepper_name }}, {{ analyst_name }}, and {{ aliquoting_name }} are trained and qualified by quality to perform the test. |  |  |  |  |  |
| # Is formula worksheet provided and verified? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  |  |
| # Does the processing method/ method set conform to the required specification? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  |  |
| # Was the sample stored appropriately? |  | # Yes No N/A |  |  |  | # Yes, the sample was stored as per client's instructions. Information is available in EagleTrax Sample Location History under {{ sample_id }} |  |  |  |  |  |
| # Material Discrepancy |  |  |  |  |  | # Lot Number |  |  |  |  | # Expiry Date |
| # Is the integrity of the sample(s) questionable? |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |
| # Are the other test results from this run under investigation? |  | # Yes No N/A (See Phase I summary) |  |  |  |  |  |  |  |  |  |
| # Appropriate /correct glassware used? |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |
| # Are reagents stored in proper condition and not expired? |  | # Yes No N/A |  |  |  | # USP <71> Process Consumables: # Information is available in EagleTrax Sample Consumables section # USP <71> Aliquoting Consumables: # See attached USP <71> data packet # Environmental Plates: # TSA and Surface Plate: see attached environmental logs # Monthly Cleaning: # H2O2 strips and IPA: See attached monthly cleaning logs |  |  |  |  | # USP <71> Process Consumables: # Information is available in EagleTrax Sample Consumables section # USP <71> Aliquoting Consumables: # See attached USP <71> data packet # Environmental Plates: # TSA and Surface Plate: see attached environmental logs # Monthly Cleaning: # H2O2 strips and IPA: See attached monthly cleaning logs |
| # Standard Used |  | # USP <71> ATP Positive Control |  |  |  | # {{ control_lot }} |  |  |  |  | # {{ control_data }} |
| # Solvent Used |  | # Not Applicable |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Is the standard the source of the root cause? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Has the analyst been instructed to keep the original test solution for possible retesting? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Equipment Discrepancy |  |  |  |  |  | # Equipment ID |  |  |  |  | # Calibration Due Date |
| # Was the instrument set up correctly? |  | # Yes No N/A |  |  |  | # E00{{ celsis_id }} |  |  |  |  | # Mar 2026 |
| # Is the clean room facility certified? |  | Yes No N/A |  |  |  | # For Processing: E00{{ cr_id }} (CR{{ cr_suit }}) # For Aliquoting: E001736 (CR114) # For Reading: E00{{ cr_id }} (CR{{ cr_suit }}) |  |  |  |  | # Jun 2026 # Dec 2026 |
| # System Performance Check (if applicable)? |  | # Yes No N/A |  |  |  | # E00{{ celsis_id }} |  |  |  |  | # Mar 2026 |
| # System Suitability RSD < 2% (PDA)? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # System Suitability RSD < 10% (QDA,CAD)? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Do the environmental monitoring results conform to required standards? |  | Yes No N/A |  |  |  | # See Phase I Summary |  |  |  |  | # See Phase I Summary |
| # Do the controls (i.e., positive, negative, etc.) conform to the required standards? |  | Yes No N/A |  |  |  | # See Phase I Summary |  |  |  |  | # See Phase I Summary |
| # Is equipment/instrument functional and calibrated/qualified/certified/PM? |  | # Yes No N/A |  |  |  | # Incubator E001034 # (Sensor E001501) # Incubator E001031 (Sensor E001505) # Incubator E001356 (Sensor E001450) # Incubator E001357 (Sensor E001449) |  |  |  |  | # Jun 2026 # Feb 2027 # Jun 2026 # Feb 2027 # Jun 2026 # Feb 2027 |
| # Are chromatogram/spectra/raw data correct? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Section C – Phase I Summary |  |  |  |  |  |  |  |  |  |  |  |
| # Has root cause been identified as a possible laboratory error? # Note: Root cause must be conclusive, verifiable, and well-documented. # Yes # If retesting is required, complete Phase II of the OOS (initiate Form 3.100.019.F02). # No, the OOS result is valid. |  |  |  |  |  |  |  |  |  |  |  |
| # Phase I Summary: # All analysts involved in the prepping, processing, aliquoting, and reading of the sample – {{ prepper_name }}, {{ analyst_name }}, and {{ aliquoting_name }} were interviewed comprehensively. Their answers are recorded throughout this document. # Upon arrival, the sample was stored in accordance with the Client’s instructions. Analyst {{ prepper_name }} verified the sample’s integrity throughout both the preparation and processing stages. No leaks or turbidity were observed at any point, verifying the integrity of the sample. # All reagents and supplies mentioned in the material section above were stored according to the suppliers’ recommendations, and their integrity was visually verified before utilization. Moreover, all reagents and supplies had valid expiration dates. The functionality of all equipment was confirmed by reviewing data generated by our comprehensive in-house continuous monitoring system. # During the preparation phase, {{ prepper_name }} disinfected the samples using acidified bleach and placed them into a pre-disinfected storage bin. On {{ test_date }}, prior to sample processing, {{ analyst_name }} performed a second disinfection with acidified bleach, allowing a minimum contact time of 10 minutes before transferring the samples into the cleanroom suites. A final disinfection step was completed immediately before the samples were introduced into the ISO 5 Biological Safety Cabinet (BSC), E00{{ bsc_id }}, located within the {{ bsc_location }}, (Suite {{ cr_suit }}{{ suit }}). All activities were conducted in accordance with SOP #2.600.059 for the USP <71> sterility testing. # {{ equipment_summary }} # On {{received_data}}, the sample vials for {{ sample_id }} were received from the Sample Submissions team and brought into the Sterile Microbiology lab. Upon arrival, each sample vial was sprayed with an acidified bleach disinfectant, placed into pre-disinfected bins, and allowed a 10-minute contact time. The secondary disinfection happened in the ISO 8 anteroom (Suite {{ cr_suit }})), where the vials were again treated with acidified bleach and provided a 10‑minute contact time before processing. Subsequently, the vials were moved into the ISO 7 cleanroom Suite {{ cr_suit }}({{ suit }}). Inside this cleanroom, the processing analyst, {{ analyst_name }}, performed a final disinfection step, allowing an additional 10-minute contact time. Once fully disinfected, the vials were transferred into the ISO 5 BSC E00{{bsc_id}}. # Once transferred into the ISO 5 BSC, the vials were placed on the disinfected working surface of the BSC E00{{bsc_id}} and aseptically opened and tested in accordance with SOP 2.600.059 (USP <71> Sterility Testing). Following testing, the media bottles were subsequently transferred into designated incubators, E001356 and E001357, to initiate incubation. # Upon completion of incubation on {{ test_date }}, both TSB & FTM bottles were disinfected and transferred to the middle ISO 7 buffer room (Suite 114A) for aliquoting step per SOP 2.600.059 (USP <71> Sterility Testing). In Suite 114A, the media bottles were disinfected one more time before transferring them to the ISO 5 BSC E001798 located in Suite 114A. In ISO 5 BSC E001798, the sample was aliquoted into assay cuvettes by analyst {{ aliquoting_name }}. After aliquoting, USP <71> Sterility Reading was performed in accordance with SOP 2.600.059 by analyst {{ aliquoting_name }}. # Following the reading, sample {{ sample_id }} was found to yield a positive reading in one of the {{positive_media}} media bottles. The average Relative Luminescence Units (RLU) from the duplicate reading tube, originating from the {{positive_media}} sample bottle, exceeded the average RLU of the {{positive_media}} negative control, confirming a positive result. The %CV from the duplicate reading tubes for the positive {{positive_media}} bottles were well within the specification (< 30%). Additionally, all Daily Controls, including the Instrument Blank, Reagent Blank, and ATP Positive Control, were within the defined specifications, each with a %CV below 30%. # Following the OOS result, the positive {{positive_media}} bottle for {{ sample_id }} was submitted for Differential Staining and Microbial Identification under {{positive_id}}, where the organisms were identified as {{positive_org}}. # The culture media utilized were within their expiry period. The negative culture media bottles for the direct inoculation method for the original culture were handled, processed, and incubated in a manner identical to that of actual samples. No microbial growth was observed in the corresponding negative control. # {{ narrative_summary }} # The analysts confirmed full compliance with cleaning procedures as outlined in SOPs 2.600.018 (Cleaning and Disinfecting Procedure for Microbiology), 2.600.059 (USP <71> Sterility Testing), and 2.600.008 (USP <71> / EP 2.6.1 Sterility Test). # Monthly cleaning and disinfection of the outermost ISO 8 Anteroom, the middle ISO 7 Buffer room, the innermost ISO 7 cleanroom, and its containing ISO 5 Biosafety Cabinets for CR {{ cr_suit }} and CR 114 was performed on {{ monthly_cleaning_date }}, as per SOP 2.600.018 (Cleaning and Disinfecting Procedure for Microbiology) During both cleaning cycles, it was documented that all H₂O₂ indicators passed. This confirms the efficient monthly cleaning of all three parts of Cleanrooms {{ cr_suit }} and 114. # {{ sample_history_paragraph }} # To assess the potential for sample-to-sample contamination contributing to the positive results, a comprehensive review was conducted of all samples processed on the same day. {{ cross_contamination_summary }} # Based on the observations outlined above, it is unlikely that the failing results were due to reagents, supplies, the cleanroom environment, the process, or analyst involvement. Consequently, the possibility of laboratory error contributing to this failure is minimal and the original result is deemed to be valid. |  |  |  |  |  |  |  |  |  |  |  |
| # Reviewer Comments: |  |  |  |  |  |  |  |  |  |  |  |
| # Acknowledgement and Closure |  |  |  |  |  |  |  |  |  |  |  |
| # Can investigation be closed? # Yes # Original result is valid. # No, initiate Form 3.100.019.F02. |  |  |  |  |  |  |  |  |  |  |  |
| # Prepared by (Print): |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |
| # Lab Manager/designee (Print): |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |
| # Approved by QA Representative (Print): # Closed: YES NO |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |

# Table 1: Information for {{ sample_id }} under investigation

| Processing Analyst | Aliquoting Analyst | Sample ID | Related Microbial ID | Media with microbial growth | Microbial ID |
| --- | --- | --- | --- | --- | --- |
| {{ analyst_name }} | {{ reader_name }} | {{ sample_id }} | {{positive_id}} | {{ positive_media}} | {{ positive_org }} |

# Table 2: Environmental Monitoring from Testing Performed on {{ process_date }}

| Environmental Monitoring (EM) Sampling Site |  | Frequency | Date (DDMMM YYYY) | Analyst (Initials) |  | Day /Week(s) |  | Observation |  | Environmental Monitoring Plate ETX ID |  | Microbial ID | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM Bracketing |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Personal (Left Touch and Right Touch) for Processing |  | Daily | {{ test_date }} | {{ analyst_initial }} |  | Date of Testing |  | {{ obs_pers_dur }} |  | {{ etx_pers_dur }} |  | {{ id_pers_dur }} | None |
| Personal (Left Touch and Right Touch) for Aliquoting |  | Daily | {{ test_date }} | {{ analyst_initial }} |  | Date of Testing |  | {{ obs_pers_dur }} |  | {{ etx_pers_dur }} |  | {{ id_pers_dur }} | None |
| Biological Safety Cabinet EM Bracketing Biological Safety Cabinet (BSC) E00{{ bsc_id }} and E001798 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 E00{{ bsc_id }} (4 locations) |  | Daily | {{ test_date }} | {{ analyst_initial }} |  | Date of Testing |  | {{ obs_surf_dur }} |  | {{ etx_surf_dur }} |  | {{ id_surf_dur }} | None |
| Surface Sampling of ISO 5 E001798 (4 locations) |  | Daily | {{ test_date }} | {{ aliquoting_initial }} |  | Date of Testing |  | {{ obs_surf_dur }} |  | {{ etx_surf_dur }} |  | {{ id_surf_dur }} | None |
| Settling Sampling of ISO 5 E00{{ bsc_id }} (2 locations) |  | Daily | {{ test_date }} | {{ analyst_initial }} |  | Date of Testing |  | {{ obs_sett_dur }} |  | {{ etx_sett_dur }} |  | {{ id_sett_dur }} | None |
| Settling Sampling of ISO 5 E001798 (2 locations) |  | Daily | {{ test_date }} | {{ aliquoting_initial }} |  | Date of Testing |  | {{ obs_sett_dur }} |  | {{ etx_sett_dur }} |  | {{ id_sett_dur }} | None |
| Weekly Active Air Sampling Bracketing – Cleanroom {{ cr_suit }} – CR{{ cr_id }} |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Environmental Monitoring (EM) Sampling Site | Frequency |  | Date (DDMMM YYYY) |  | Analyst (Initials) |  | Observation |  | Environmental Monitoring Plate ETX ID |  | Microbial ID |  | Notes |
| Active Air Sampling of Cleanrooms | Weekly |  | {{ date_of_weekly }} |  | {{ weekly_initial }} |  | {{ obs_air_wk_of }} |  | {{ etx_air_wk_of }} |  | {{ id_air_wk_of }} |  | None |
| Surface Sampling of Anteroom and Cleanroom Bracketing - Cleanroom {{ cr_suit }} – CR{{ cr_id }} |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms | Weekly |  | {{ date_of_weekly }} |  | {{ weekly_initial }} |  | {{ obs_room_wk_of }} |  | {{ etx_room_wk_of }} |  | {{ id_room_wk_of }} |  | None |