Table 1: Information for ETX-260624-0251 (BUD-T90) under investigation

| Processing Analyst | Reading Analyst | Sample ID | Related Microbial ID | Media with microbial growth | Microbial ID |
| --- | --- | --- | --- | --- | --- |
| Guanchen Li | Andrew Carrillo | ETX-260624-0251 (BUD-T90) | ETX-261012-0367 | TSB | Bacillus megaterium |

Table 2: Environmental Monitoring from Processing Performed on 23Sep26

| Environmental Monitoring (EM) Sampling Site | Frequency | Date (DDMMMYY) | Analyst | Day /Week(s) | Observation | EM Plate ETX Number | Microbial ID | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM |  |  |  |  |  |  |  |  |
| Personnel (Left Touch and Right Touch) | Daily | 22Sep26 | GL | Date Before Testing | No Growth | Not Applicable | Not Applicable | None |
| Personnel (Left Touch and Right Touch) | Daily | 23Sep26 | GL | Date of Testing | No Growth | Not Applicable | Not Applicable | None |
| Personnel (Left Touch and Right Touch) | Daily | 24Sep26 | GL | Date After Testing | No Growth | Not Applicable | Not Applicable | None |
| Biological Safety Cabinet EM for E001314 |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 (4 locations) | Daily | 22Sep26 | AOD | Date Before Testing | No Growth | Not Applicable | Not Applicable | None |
| Surface Sampling of ISO 5 (4 locations) | Daily | 23Sep26 | GL | Date of Testing | No Growth | Not Applicable | Not Applicable | None |
| Surface Sampling of ISO 5 (4 locations) | Daily | 24Sep26 | AOD | Date After Testing | No Growth | Not Applicable | Not Applicable | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 22Sep26 | AOD | Date Before Testing | No Growth | Not Applicable | Not Applicable | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 23Sep26 | GL | Date of Testing | No Growth | Not Applicable | Not Applicable | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 24Sep26 | AOD | Date After Testing | No Growth | Not Applicable | Not Applicable | None |
| Weekly Active Air Sampling of Cleanroom Suite 115 (E001737) with Processing BSC |  |  |  |  |  |  |  |  |
| Active Air Sampling of Cleanrooms | Weekly | 23Sep26 | SMO | Week of Testing | 6 CFUs (ISO 8 115) | ETX-261001-0625 | Micrococcus luteus Kocuria indica Kocuria rhizophila Pending FunITS | None |
| Weekly Surface Sampling of Cleanroom Suite 115 (E001737) with Processing BSC |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms | Weekly | 23Sep26 | SMO | Week of Testing | No Growth | Not Applicable | Not Applicable | None |