Table 1: Information for ETX-260929-0931 under investigation

| Processing Analyst | Reading Analyst | Sample ID | Events | Confirmed Microbial Events | Morphology Description |
| --- | --- | --- | --- | --- | --- |
| Ellen Bennett | Varsha Subramanian | ETX-260929-0931 | 15 | 4 | Cocci shaped morphology |

Table 2: Environmental Monitoring from Testing Performed on 01Oct26

| Environmental Monitoring (EM) Sampling Site | Frequency | Date (DDMMMYY) | Analyst | Day /Week(s) | Observation | EM Plate ETX Number | Microbial ID | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM |  |  |  |  |  |  |  |  |
| Personal (Left Touch and Right Touch) | Daily | 01Oct26 | ELB | Date of Testing (Scan) | No Growth | Not Applicable | Not Applicable | None |
| Personal (Left Touch and Right Touch) | Daily | 01Oct26 | KSM | Date of Testing (Scan C/O) | No Growth | Not Applicable | Not Applicable | None |
| Biological Safety Cabinet EM for E001938 and E001937 |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 BSC E001938 (4 Locations) | Daily | 01Oct26 | ELB | Date of Testing (Scan) | No Growth | Not Applicable | Not Applicable | None |
| Surface Sampling of ISO 5 BSC E001937 (4 Locations) | Daily | 01Oct26 | KSM | Date of Testing (Scan C/O) | No Growth | Not Applicable | Not Applicable | None |
| Settling Sampling of ISO 5 BSC E001938 | Daily | 01Oct26 | ELB | Date of Testing (Scan) | No Growth | Not Applicable | Not Applicable | None |
| Settling Sampling of ISO 5 BSC E001937 | Daily | 01Oct26 | KSM | Date of Testing (Scan C/O) | No Growth | Not Applicable | Not Applicable | None |
| Weekly Active Air Sampling of Cleanroom 145, 144, 143, 142 (L-Suite) with Processing BSCs |  |  |  |  |  |  |  |  |
| Active Air Sampling of Cleanrooms | Weekly | 01Oct26 | SMO | Week of Testing | 1 CFU (ISO 7 144 Sec II) | ETX-261012-0430 | Staphylococcus epidermidis | None |
|  |  |  |  |  | 5 CFUs (ISO 8 142) | ETX-261012-0433 | Staphylococcus epidermidis Pseudomonas parafulva Bacillus safensis Corynebacterium xerosis Corynebacterium ureicelerivorans |  |
| Weekly Surface Sampling of Cleanroom 145, 144, 143, 142 (L-Suite) with Processing BSCs |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms | Weekly | 01Oct26 | SMO | Week of Testing | No Growth | Not Applicable | Not Applicable | None |