Phase I

| # Section A – Test Details |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| # Initiator: | # Alex Saravia # (Written by: Qiyue Chen) | # Test Date: |  | # 31Aug26 |  |  |  | # Date Initiated: |  | # 07Sep26 |  |
| # Date of Incident: | # 07Sep26 | # Test Name: |  | # USP <71> Sterility Test |  |  |  | # Sample ID: |  | # ETX-260806-0702 (BUD-T30) # Hillstone Pharmacy (E17047) |  |
| # Name of Analyst who Performed the Test: | # Prepper: # Andrew Carrillo (AC) # Processor: Alex Saravia (ES) # Subculture Processor: () # Reading Analyst: # Andrew Carrillo (AC) | # Sample / Active Name |  | # TB-500 10mg/ml |  |  |  | # Lot #: |  | # 24JUL2026-059 |  |
|  |  | # Dosage Form |  | # Liquid |  |  |  |  |  |  |  |
| # Description of # Incident: | # On 07Sep26, sample ETX-260806-0702 (BUD-T30) was found positive for viable microorganisms after USP <71> sterility testing. |  |  |  | # Client Care notified of OOS # Sterility and Endotoxin test related OOS; Client must be notified within 24 hours of OOS |  |  |  |  |  |  |
|  | # SOP / Test Method #: # MICRO-SOP-5 (19) | # Effective Date: # 24Jul26 |  |  |  |  | # SOP / Test Method Rev: # Rev: 19 |  | # Limits/Specification: # Pass/Fail |  |  |
| # Manager notified | # Manager/designee name: # Kathan Parikh | # Date: # 07Sep26 |  |  |  |  | # Quality Assurance Department Notified: |  | # Yes No |  |  |
| # Section B – Investigation and Root Cause Analysis |  |  |  |  |  |  |  |  |  |  |  |
| # Personnel Discrepancy |  |  |  |  |  | # Comments |  |  |  |  |  |
| # Analyst interviewed? |  | # Yes No N/A |  |  |  | # Yes, analysts Andrew Carrillo, Alex Saravia, , and Andrew Carrillo were interviewed comprehensively. |  |  |  |  |  |
| # Correct samples analyzed? |  | # Yes No N/A |  |  |  | # Yes, sample ID: ETX-260806-0702 (BUD-T30) |  |  |  |  |  |
| # Correct SOP / Test Method used/followed? |  | # Yes No N/A |  |  |  | # Yes, as per MICRO-SOP-5 |  |  |  |  |  |
| # Correct sample collection/technique? |  | # Yes No N/A |  |  |  | # Yes, as per MICRO-SOP-5 |  |  |  |  |  |
| # Are the Test Result Records Attached & Calculations Verified and Correct? |  | # Yes No N/A |  |  |  | # Yes, Information is available in EagleTrax under ETX-260806-0702 (BUD-T30) |  |  |  |  |  |
| # Is the analyst qualified to perform the test? |  | # Yes No N/A |  |  |  | # Yes, all analysts are trained and qualified by quality to perform the test. |  |  |  |  |  |
| # Is formula worksheet provided and verified? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  |  |
| # Does the processing method/ method set conform to the required specification? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  |  |
| # Was the sample stored appropriately? |  | # Yes No N/A |  |  |  | # Yes, the sample was stored as per client's instructions. Information is available in EagleTrax Sample Location History under ETX-260806-0702 (BUD-T30) |  |  |  |  |  |
| # Material Discrepancy |  |  |  |  |  | # Lot Number |  |  |  |  | # Expiry Date |
| # Is the integrity of the sample(s) questionable? |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |
| # Are the other test results from this run under investigation? |  | # Yes No N/A (See Phase I summary) |  |  |  |  |  |  |  |  |  |
| # Appropriate /correct glassware used? |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |
| # Are reagents stored in proper condition and not expired? |  | # Yes No N/A |  |  |  | # Process Consumables: # Information is available in EagleTrax Sample Consumables section # Environmental Plates: # TSA and Surface Plate: see attached environmental logs # Monthly Cleaning: # H2O2 strips and IPA: See attached monthly cleaning logs |  |  |  |  | # Process Consumables: # Information is available in EagleTrax Sample Consumables section # Environmental Plates: # TSA and Surface Plate: see attached environmental logs # Monthly Cleaning: # H2O2 strips and IPA: See attached monthly cleaning logs |
| # Standard Used |  | # Not Applicable |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Solvent Used |  | # Not Applicable |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Is the standard the source of the root cause? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Has the analyst been instructed to keep the original test solution for possible retesting? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Equipment Discrepancy |  |  |  |  |  | # Equipment ID |  |  |  |  | # Calibration Due Date |
| # Was the instrument set up correctly? |  | # Yes No N/A |  |  |  | # ISO 5 BSC E001316 |  |  |  |  | # Mar 2026 |
| # Is the clean room facility certified? |  | Yes No N/A |  |  |  | # E001736 (CR114) |  |  |  |  | # Jun 2026 # Dec 2026 |
| # System Performance Check (if applicable)? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Mar 2026 |
| # System Suitability RSD < 2% (PDA)? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # System Suitability RSD < 10% (QDA,CAD)? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Do the environmental monitoring results conform to required standards? |  | Yes No N/A |  |  |  | # See Phase I Summary |  |  |  |  | # See Phase I Summary |
| # Do the controls (i.e., positive, negative, etc.) conform to the required standards? |  | Yes No N/A |  |  |  | # See Phase I Summary |  |  |  |  | # See Phase I Summary |
| # Is equipment/instrument functional and calibrated/qualified/certified/PM? |  | # Yes No N/A |  |  |  | # Incubator E001034 # (Sensor E001501) # Incubator E001031 (Sensor E001505) # Incubator E001356 (Sensor E001450) # Incubator E001357 (Sensor E001449) |  |  |  |  | # Jun 2026 # Feb 2027 # Jun 2026 # Feb 2027 # Jun 2026 # Feb 2027 |
| # Are chromatogram/spectra/raw data correct? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Section C – Phase I Summary |  |  |  |  |  |  |  |  |  |  |  |
| # Has root cause been identified as a possible laboratory error? # Note: Root cause must be conclusive, verifiable, and well-documented. # Yes # If retesting is required, complete Phase II of the OOS (initiate Form 3.100.019.F02). # No, the OOS result is valid. |  |  |  |  |  |  |  |  |  |  |  |
| # Phase I Summary: All analysts involved in the prepping, processing, and reading of the sample – Andrew Carrillo and Alex Saravia, were interviewed comprehensively. Their answers are recorded throughout this document. Upon arrival, the sample was stored in accordance with the Client’s instructions. Analysts - Andrew Carrillo and Alex Saravia - verified the integrity of the sample throughout both the preparation and processing stages. No leaks or turbidity were observed at any point, verifying the integrity of the sample. All reagents and supplies mentioned in the material section above were stored according to the suppliers’ recommendations, and their integrity was visually verified before utilization. Moreover, all reagents and supplies had valid expiration dates. The functionality of all equipment was confirmed by reviewing data generated by our comprehensive in-house continuous monitoring system. The cleanroom suite (CR 114) used for processing procedures comprises three interconnected sections: the innermost ISO 7 cleanroom (114B), which connects to the middle ISO 7 buffer room (114A), and then to the outermost ISO 8 anteroom (114). A positive air pressure system is maintained throughout the suite to ensure controlled, unidirectional airflow from 114B through 114A and into 114. The ISO 5 BSC E001316, located in the innermost ISO 7 room, (114B), was used for sample processing steps. It was thoroughly cleaned and disinfected prior to each procedure in accordance with MICRO SOP-9 (Cleaning and Disinfecting Procedure for Microbiology). Additionally, BSC E001316 was certified and approved by both the Engineering and Quality Assurance teams. Sample processing was conducted in the ISO 5 BSC E001316 in the innermost ISO 7 room, (114B) by Alex Saravia on 31 Aug 2026 as per MICRO-SOP-5 (USP 71> / EP 2.6.1 Sterility Test). On 28 Aug 2026, the sample vials for ETX-260806-0702-T30 were received from the Sample Submissions team and brought into the Sterile Microbiology lab. Upon arrival, each sample vial was sprayed with an acidified bleach disinfectant, placed into pre-disinfected bins, and allowed a 10-minute contact time. The secondary disinfection happened in the ISO 8 anteroom (114), where the vials were again treated with acidified bleach and provided a 10-minute contact time before processing. Subsequently, the vials were moved into innermost ISO 7 cleanroom (114B). Inside this cleanroom, the processing analyst, Alex Saravia, performed a final disinfection step, allowing an additional 10-minute contact time. Once fully disinfected, the vials were transferred into the ISO 5 BSC E001316. Inside the BSC, each container was aseptically opened, ensuring minimal exposure to the environment. The contents of 10 vials (20 mL per media type) were dispensed directly from the container into sterile canisters placed on a vacuum manifold for membrane filtration under method suitability ETX-260623-0160. The sample was processed and tested as per MICRO-SOP-5 (USP 71> / EP 2.6.1 Sterility Test). Media bottles were transferred into incubators E001356 and E001357 to begin incubation. On 07 Sep 2026 (Day 7 of incubation), microbial growth was observed in one TSB (Tryptic Soy Broth) media bottle by reading analyst, Andrew Carrillo, and confirmed by Microbiology Lab Supervisor, Robin Seymour. The positive TSB bottle for ETX-260806-0702-T30 was submitted for Differential Staining and Microbial Identification under ETX-260907-0017. Definitive identification identified the isolate as Paenibacillus lautus, Gram (+) rods. ISO 5 BSC Environmental Monitoring Results Evaluation: After reviewing the Environmental Monitoring (EM) results for the relevant testing period, no microbial growth was detected on personnel monitoring plates, settling plates, or ISO 5 BSC surface sampling plates on the date of testing, the preceding sampling date, or the subsequent sampling date. Cleanroom Environmental Monitoring Results Evaluation: No microbial growth was recovered from the weekly active surface sampling plate collected on 04 Sep 2026. However, 2 CFUs (ETX-260914-0487) were recovered from the weekly active air sampling of ISO 8 cleanroom 114. The organism was identified to be of Gram-positive coccobacilli and Gram-positive rod morphology. It is important to note that all sample processing activities were performed within the validated ISO 5 BSC E001316 located in the innermost ISO 7 cleanroom (114B). The test samples do not come into contact with ambient ISO 8 air, as samples are transferred in disinfected, lidded containers through the layered cleanroom suites. Furthermore, the colony recovery occurred in the lower-classified ISO 8 anteroom environment, which is physically segregated from the ISO 5 processing zone. In addition, considering the facility's layered disinfection program, controlled unidirectional airflow, and cleanroom design, it is unlikely that the microbial contaminant detected in the ISO 8 environment impacted on the sample processed within the ISO 5 area of the ISO 7 cleanroom (114B). The absence of contamination on analyst glove plates and work-surface environmental monitoring plates collected on the date of testing indicates that there was no viable transfer of contaminants to the ISO 5 biological safety cabinet (BSC). Additionally, the lack of contamination in other samples processed during the same testing session further demonstrates that the testing environment was operating in optimal condition. Therefore, considering that no viable organisms were recovered from the ISO 5 environment, personnel monitoring results were acceptable on the date of testing, and the environmental isolates were not recovered from locations or surfaces directly involved in sample preparation or testing, the environmental isolate is not considered a credible source of the contamination detected in the test sample. The analysts confirmed full compliance with cleaning procedures as outlined in MICRO SOP-9 (Cleaning and Disinfecting Procedure for Microbiology) and MICRO-SOP-5 (USP 71> / EP 2.6.1 Sterility Test). Monthly cleaning and disinfection of the outermost ISO 8 Anteroom, the middle ISO 7 Buffer room, the innermost ISO 7 cleanroom, and its containing ISO 5 BSCs for Cleanroom suite CR114 was performed on 30 Aug 2026, as per MICRO-SOP-9 (Cleaning and Disinfecting Procedure for Microbiology). During both cleaning cycles, it was documented that all H2O2 indicators passed. This confirms the efficient monthly cleaning of all three parts of Cleanroom suite CR114. Analyzing a 6-month sample history for Hillstone Pharmacy (E17047), this specific analyte "TB-500 10mg/ml" has had no prior failures using USP 71> / EP 2.6.1 Sterility Test during this period. This sample was submitted as part of a Beyond-Use Date (BUD) stability study for lot 24JUL2026-059. The baseline T0 time point (tested on 06 Aug 2026 by analyst GL) yielded valid passing results for USP 71> sterility. To assess the potential for sample-to-sample contamination contributing to the positive results, a comprehensive review was conducted of all samples processed on the same day (31 Aug 2026). Sample ETX-260806-0702-T30 was the 4th sample processed in the batch. All other samples processed by the analyst and other analysts that day tested negative. These findings suggest that cross-contamination between samples is highly unlikely. Based on the observations outlined above, it is unlikely that the failing results were due to reagents, supplies, the cleanroom environment, the process, or analyst involvement. Consequently, the possibility of laboratory error contributing to this failure is minimal and the original result is deemed to be valid. |  |  |  |  |  |  |  |  |  |  |  |
| # Reviewer Comments: |  |  |  |  |  |  |  |  |  |  |  |
| # Acknowledgement and Closure |  |  |  |  |  |  |  |  |  |  |  |
| # Can investigation be closed? # Yes # Original result is valid. # No, initiate Form 3.100.019.F02. |  |  |  |  |  |  |  |  |  |  |  |
| # Prepared by (Print): |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |
| # Lab Manager/designee (Print): |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |
| # Approved by QA Representative (Print): # Closed: YES NO |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |

# Table 1: Information for ETX-260806-0702 (BUD-T30) under investigation

| Processing Analyst | Reading Analyst | Sample ID | Related Microbial ID | Media with Microbial growth | Morphology Description |
| --- | --- | --- | --- | --- | --- |
| Alex Saravia | Andrew Carrillo | ETX-260806-0702-T30 | ETX-260907-0017 | 1 x 100 mL TSB | Paenibacillus lautus (Gram (+) rods) |

# Table 2: Environmental Monitoring for Analyst & Cleanroom

| Environmental Monitoring (EM) Sampling Site | Frequency | Date (DDMMMYYYY) | Analyst (Initials) | Day /Week(s) | Observation | Environmental Monitoring Plate ETX ID | Microbial ID | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM Bracketing 31Aug2026 |  |  |  |  |  |  |  |  |
| Personal (Left Touch and Right Touch) | Daily | 28Aug2026 | ES | Date Before Testing | No growth | N/A | N/A | None |
| Personal (Left Touch and Right Touch) | Daily | 31Aug2026 | ES | Date of Testing | No growth | N/A | N/A | None |
| Personal (Left Touch and Right Touch) | Daily | 01Sep2026 | ES | Date After Testing | No growth | N/A | N/A | None |
| Biological Safety Cabinet EM Bracketing Biological Safety Cabinet (BSC) E001316 |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 (4 locations) | Daily | 28Aug2026 | ES | Date Before Testing | No growth | N/A | N/A | None |
| Surface Sampling of ISO 5 (4 locations) | Daily | 31Aug2026 | ES | Date of Testing | No growth | N/A | N/A | None |
| Surface Sampling of ISO 5 (4 locations) | Daily | 01Sep2026 | ES | Date After Testing | No growth | N/A | N/A | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 28Aug2026 | ES | Date Before Testing | No growth | N/A | N/A | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 31Aug2026 | ES | Date of Testing | No growth | N/A | N/A | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 01Sep2026 | ES | Date After Testing | No growth | N/A | N/A | None |
| Weekly Active Air Sampling of CR 114 (E001736) |  |  |  |  |  |  |  |  |
| Active Air Sampling of Cleanrooms | Weekly | 04Sep2026 | SMO | Week of Testing | 2 CFUs (ISO 8 114) | ETX-260914-0487 | Gram (+) rods and Gram (+) coccobacilli | None |
| Surface Sampling of Anteroom and Cleanroom of CR 114 (E001736) |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms | Weekly | 04Sep2026 | SMO | Week of Testing | No growth | N/A | N/A | None |