Phase I

| # Section A – Test Details |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| # Initiator: | # Simin Mohammad | # Test Date: |  | # 08-May-2026 |  |  |  | # Date Initiated: |  | 18-May-2026 |  |
| # Date of Incident: | 18-May-2026 | # Test Name: |  | Environmental Monitoring |  |  |  | # Sample ID: |  | # ETX-260518-0250 |  |
| # Name of Analyst who Performed the Test: | # Simin Mohammad (Weekly Active Air Sampling Plate Setup) Maraya Chukwumerije (Weekly Active Air Sampling Plate Reader) Sophia Santamaria (Weekly Active Air Sampling Plate Reader) | # Sample / Active Name |  | # EM SMO 115B Air 08MAY2026 |  |  |  | # Lot #: |  | # EM SMO 115B Air 08MAY2026 |  |
|  |  | # Dosage Form |  | # Plate |  |  |  |  |  |  |  |
| # Description of # Incident: | # The CFU count for the environmental monitoring plate exceeded the action level. |  |  |  | # Client Care notified of OOS # Sterility and Endotoxin test related OOS; Client must be notified within 24 hours of OOS |  |  |  |  |  |  |
|  | SOP / Test Method #: MICRO-SOP-2 | Effective Date: 23-Jul-2026 |  |  |  |  | SOP / Test Method Rev: 16 |  | # Limits/Specification: # <1 event |  |  |
| # Manager notified | # Manager/designee name: # Kathan Parikh | Date: 18-May-2026 |  |  |  |  | # Quality Assurance Department Notified: |  | # Yes No |  |  |
| # Section B – Investigation and Root Cause Analysis |  |  |  |  |  |  |  |  |  |  |  |
| # Personnel Discrepancy |  |  |  |  |  | # Comments |  |  |  |  |  |
| # Analyst interviewed? |  | # Yes No N/A |  |  |  | # Yes, analysts Simin Mohammad, Maraya Chukwumerije, and Sophia Santamaria were comprehensively interviewed. |  |  |  |  |  |
| # Correct samples analyzed? |  | # Yes No N/A |  |  |  | # Yes, as per MICRO-SOP-2 |  |  |  |  |  |
| # Correct SOP / Test Method used/followed? |  | # Yes No N/A |  |  |  | Yes, as per MICRO-SOP-2 |  |  |  |  |  |
| # Correct sample collection/technique? |  | # Yes No N/A |  |  |  | Yes, as per MICRO-SOP-2 |  |  |  |  |  |
| # Are the Test Result Records Attached & Calculations Verified and Correct? |  | # Yes No N/A |  |  |  | # Yes, Information is available in EagleTrax under ETX-260518-0250 |  |  |  |  |  |
| # Is the analyst qualified to perform the test? |  | # Yes No N/A |  |  |  | Yes, the analysts are trained and qualified by quality to perform the test |  |  |  |  |  |
| # Is formula worksheet provided and verified? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  |  |
| # Does the processing method/ method set conform to the required specification? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  |  |
| # Was the sample stored appropriately? |  | # Yes No N/A |  |  |  | Yes, as per MICRO-SOP-2 |  |  |  |  |  |
| # Material Discrepancy |  |  |  |  |  | # Lot Number |  |  |  |  | # Expiry Date |
| # Is the integrity of the sample(s) questionable? |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |
| # Are the other test results from this run under investigation? |  | # Yes No N/A (See Phase I summary) |  |  |  |  |  |  |  |  |  |
| # Appropriate /correct glassware used? |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |
| # Are reagents stored in proper condition and not expired? |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |
| # Standard Used |  |  |  |  |  |  |  |  |  |  |  |
| # Solvent Used |  | # Not Applicable |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Is the standard the source of the root cause? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Has the analyst been instructed to keep the original test solution for possible retesting? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Equipment Discrepancy |  |  |  |  |  | # Equipment ID |  |  |  |  | # Calibration Due Date |
| # Was the instrument set up correctly? |  | # Yes No N/A |  |  |  | Please see below |  |  |  |  | Please see below |
| # Is the clean room facility certified? |  | Yes No N/A |  |  |  | CR115 (Sensor E001737) |  |  |  |  | December 2026 |
| # System Performance Check (if applicable)? |  | # Yes No N/A |  |  |  | N/A |  |  |  |  | N/A |
| # System Suitability RSD < 2% (PDA)? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # System Suitability RSD < 10% (QDA,CAD)? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Do the environmental monitoring results conform to required standards? |  | Yes No N/A |  |  |  | # See Phase I Summary |  |  |  |  | # See Phase I Summary |
| # Do the controls (i.e., positive, negative, etc.) conform to the required standards? |  | Yes No N/A |  |  |  | # See Phase I Summary |  |  |  |  | # See Phase I Summary |
| # Is equipment/instrument functional and calibrated/qualified/certified/PM? |  | # Yes No N/A |  |  |  | # Incubator E001034 # (Sensor E001501) # Incubator E001031 (Sensor E001505) |  |  |  |  | # Aug26 # Feb26 # Aug26 # Feb26 |
| # Are chromatogram/spectra/raw data correct? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Section C – Phase I Summary |  |  |  |  |  |  |  |  |  |  |  |
| # Has root cause been identified as a possible laboratory error? # Note: Root cause must be conclusive, verifiable, and well-documented. # Yes # If retesting is required, complete Phase II of the OOS (initiate Form 3.100.019.F02). # No, the OOS result is valid. |  |  |  |  |  |  |  |  |  |  |  |
| Phase I Summary: # The analyst involved in the Active Air Sampling plate setup, Simin Mohammad, and the analysts involved in reading the plate, Maraya Chukwumerije and Sophia Santamaria, were interviewed comprehensively. Their answers are recorded throughout this document. The EM plates were stored in compliance with the supplier's recommendations, and their integrity was visually inspected prior to use. Furthermore, the plates were confirmed to be within their valid expiration dates. All the supplies were thoroughly disinfected according to MICRO-SOP-9 (Cleaning and Disinfecting Procedure for Microbiology). The functionality of both incubators was verified through a review of data obtained from our comprehensive in-house continuous monitoring system. Active Air Sampling was performed by analyst Simin Mohammad during weekly environmental monitoring processing in Suite 115 (ISO 8), Suite 115A (ISO 7), and Suite 115B (ISO 7) on 08 May 2026 as per MICRO-SOP-2 (Environmental Monitoring of the Cleanroom Facility). The plates were initially incubated at a temperature of 30–35°C in incubator E001031 for a minimum duration of 48 hours, commencing on 08 May 2026. Following completion of a minimum of 48 hours of incubation on 11 May 2026, the plates were further incubated for a minimum of 5 days at 20–25°C in incubator E001034, with the incubation concluding on 18 May 2026. Please see Table 1 for detailed information on the observations during respective incubations. Based on the observations in Table 1, since the CFU count exceeded the action level for one of the Active Air Sampling plates (115B), the plate was submitted for Microbial Identification under ETX-260518-0250. The colonies were identified as Penicillium decumbens, Cladosporium tenuissimum, Cladosporium langeronii, and Cladosporium halotolerans. To observe if the organisms identified were transient in nature or recurring, weekly environmental monitoring plates for the clean room were bracketed to include the week before testing, the week of testing, and the week after testing as detailed in Table 2 (please see attached). Environmental Monitoring Summary: Weekly environmental sampling for the previous week showed 1 CFU for 115 ISO 8 identified as Micrococcus luteus. Week of testing showed 1 CFU in 115A ISO 7 and 11 CFU in 115B ISO 7 identified as Staphylococcus aureus, Penicillium decumbens, Cladosporium tenuissimum, Cladosporium langeronii, and Cladosporium halotolerans. Lastly, the following week of testing showed 1 CFU in 115 ISO 8, 11 CFU in 115B ISO 7, 1 CFU on table in ISO 8 115, and 1 CFU on the cart in 115 ISO 8 identified as Gram (+) cocci, Candida orthopsilosis, Budding yeast, Gram (+) rods, and Staphylococcus capitis. During the interview with the analyst, they indicated that no obvious abnormalities or deviations in the testing procedure were observed. All the samples were thoroughly disinfected prior to testing. Moreover, the CR115 was thoroughly cleaned and prepared before initiating the testing as per MICRO-SOP-2 (Environmental Monitoring of the Cleanroom Facility) and MICRO-SOP-9 (Cleaning and Disinfecting Procedure for Microbiology). Monthly cleaning and disinfection of the outermost ISO 8 Anteroom (Suite 115), the middle ISO 7 Buffer room (Suite 115A), the innermost ISO 7 clean room (Suite 115B), and its containing ISO 5 Biosafety Cabinets were performed on 26-Apr-2026 by analysts Tamiru Kotisso and Cuong Du as per MICRO-SOP-9 (Cleaning and Disinfecting Procedure for Microbiology). It was documented that all H2O2 indicators passed. This confirms the efficient monthly cleaning of all three suites 115, 115A, and 115B. Additionally, cleaning and disinfecting was performed both prior to and after the testing process as per MICRO-SOP-9. # Based on the findings outlined in the preceding sections, the Out-Of-Specification (OOS) result observed for the Environmental Monitoring (EM) Active Air Sampling plate may be attributed to a potential analyst error. It is to be noted that the growth observed on the weekly active air and weekly surface sampling plates for the day of testing do not follow a trend, indicating that the contamination was transient in nature and that routine daily disinfection procedures were effective in eliminating the contamination. Furthermore, no trend was observed in the clean rooms previous weekly EM data, therefore, no preventive and corrective actions are deemed necessary at this time. |  |  |  |  |  |  |  |  |  |  |  |
| # Reviewer Comments: |  |  |  |  |  |  |  |  |  |  |  |
| # Acknowledgement and Closure |  |  |  |  |  |  |  |  |  |  |  |
| # Can investigation be closed? # Yes # Original result is valid. # No, initiate Form 3.100.019.F02. |  |  |  |  |  |  |  |  |  |  |  |
| # Prepared by (Print): |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |
| # Lab Manager/designee (Print): |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |
| # Approved by QA Representative (Print): # Closed: YES NO |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |

# Table 1: Information for ETX-260518-0250 under investigation

# Table 2: Environmental Monitoring from Testing Performed on 08-May-2026

# Table 3: In Trend of Past OOS Results for Eagle Analytical Internal EM for analyte "EM SMO 115B Air 08MAY2026"

Table 1: Read Dates & Incubation Observation

Table 2: Environmental Monitoring for Analyst & Cleanroom

| ETX Submission ID | Set-up Analyst & Date | Sampling Site & Location | Plate Reading Analyst (≥ 48H) | CFUs Observed after 48 Hour Incubation at 30–35°C (E001031) | Plate Reading Analyst (NLT 5 days) | CFUs Observed after NLT 5-day Incubation at 20–25°C (E001034) | Microbial Identification |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ETX-260518-0250 | SMO 08-May-2026 |  |  |  |  |  |  |

| Environmental Monitoring (EM) Sampling Site | Frequency | Date (DDMMMYYYY) | Analyst (Initials) | Day /Week(s) | Observation | Environmental Monitoring Plate ETX ID | Microbial ID | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM Bracketing |  |  |  |  |  |  |  |  |
| Personal (Left Touch and Right Touch) | Daily |  | SMO | Date Before Testing |  |  |  | None |
| Personal (Left Touch and Right Touch) | Daily | 08-May-2026 | SMO | Date of Testing |  |  |  | None |
| Personal (Left Touch and Right Touch) | Daily |  | SMO | Date After Testing |  |  |  | None |
| Biological Safety Cabinet EM Bracketing Biological Safety Cabinet (BSC) |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  |  | Date Before Testing |  |  |  | None |
| Surface Sampling of ISO 5 (4 locations) | Daily | 08-May-2026 |  | Date of Testing |  |  |  | None |
| Surface Sampling of ISO 5 (4 locations) | Daily |  |  | Date After Testing |  |  |  | None |
| Settling Sampling of ISO 5 (2 locations) | Daily |  |  | Date Before Testing |  |  |  | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 08-May-2026 |  | Date of Testing |  |  |  | None |
| Settling Sampling of ISO 5 (2 locations) | Daily |  |  | Date After Testing |  |  |  | None |
| Weekly Active Air Sampling Bracketing |  |  |  |  |  |  |  |  |
| Active Air Sampling of Cleanrooms | Weekly |  |  | Week (On or After Testing Date) |  |  |  | None |
| Surface Sampling of Anteroom and Cleanroom Bracketing |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms | Weekly |  |  | Week (On or After Testing Date) |  |  |  | None |