| EAGLE | Category: Form Corporate Name: Nonconformance Report Form | ETX ID: 3.100.007.F01 | Effective: 10-JUL-2026 |
| --- | --- | --- | --- |
| CORP-FORM-5 v9.1 | Location: Houston (Corporate) | Printed from https://app.zenqms.com | NCR#: N 26487 |

IDENTIFICATION

| Initiator: | Qiyue Chen | Date: | 18-Sep-2026 |
| --- | --- | --- | --- |
| Department: | Microbiology | Equipment ID #: | E001230 (Scan RDI) |
| Sample ID #: | ETX-260915-0653 ETX-260916-0044 ETX-260916-0374 | Sample Name / Description: | Semaglutide/Cyanocobalamin Tirzepatide 22mg/mL-Pyridoxine HCl 4mg/mL SEMAGLUTIDE/ B12 1.2 MG/ 500 MCG/ML Inj |
| Lot #: | 0216202602 22666 LG342010613 | Client Info: | ANG Labs (E73555) Southend Pharmacy (E19207) Optimal Balance Pharmacy (E19193) |

NONCONFORMITY DESCRIPTION (Who, What, When, Where, Why, & How)

On 17-Sep-2026, during routine Scan RDI sterility testing operations, an unexpected software crash occurred on the Scan RDI instrument (Equipment ID: E001230) while running session '17 Sep 2026 - 3'. Third-shift analyst JOC immediately notified Project Microbiologist QYC regarding the incident and software failure. The following day, on 18-Sep-2026, QYC attempted to re-open the application and recover the testing session data; however, the instrument repeatedly displayed error messages stating 'Unable to Create New Session Unknown Error' and 'Unable to Load Session Unknown Error', preventing access to the session. On 21-Sep-2026, QYC informed Eagle IT Coordinator KT (Kevin Torres) and together they attempted to restore and retrieve the session data. Following technical review, KT confirmed that the session file was corrupted, exporting as only 7 KB, and the corresponding raw acquisition dataset within the D:\ directory was completely empty and unrecoverable. A total of three (3) sterility test sample submissions were actively processed in this corrupted session and impacted by data loss: • ETX-260915-0653 ¦ Client: ANG Labs (E73555) ¦ Sample: Semaglutide/Cyanocobalamin ¦ Lot#: 0216202602 • ETX-260916-0044 ¦ Client: Southend Pharmacy (E19207) ¦ Sample: Tirzepatide 22mg/mL-Pyridoxine HCl 4mg/mL (1 mL) ¦ Lot#: 22666 • ETX-260916-0374 ¦ Client: Optimal Balance Pharmacy (E19193) ¦ Sample: SEMAGLUTIDE/ B12 1.2 MG/ 500 MCG/ML (2 ML) Inj ¦ Lot#: LG342010613 No valid test results could be generated or verified for these submissions from the corrupted run. Work Order WO-260470 was promptly issued to Engineering/IT to formally investigate the software failure and initiate technical remediation.

CAUSE(S) ANALYSIS

Description of Cause(s): Scan RDI software crash and session file corruption on instrument E001230, resulting in unrecoverable raw acquisition data in the D:\ directory.

Category: [X] Equipment/System [ ] Process/Method [ ] Personnel [ ] External Phenomena [ ] Other

INTERIM CORRECTIVE ACTION(S)

1. Work Order WO-260470 was issued to Engineering and IT to investigate and resolve the software crash, clear corrupted temporary files, and evaluate database integrity on instrument E001230. 2. Instrument E001230 was taken out of service for live sample testing pending completion and sign-off of WO-260470. 3. Client Care was instructed to request additional sample vials for retesting from ANG Labs (ETX-260915-0653) and Southend Pharmacy (ETX-260916-0044). Additional sample for ETX-260916-0374 (Optimal Balance Pharmacy) was received and is in-process with the third shift for retesting on an alternate qualified Scan RDI unit.

QUALITY EVALUATION & CLOSING

| Was action effective in addressing root cause? | [X] Yes [ ] No [ ] N/A |
| --- | --- |
| Additional Comments / Risk Assessment: | Impact is isolated to the 3 sample submissions processed in the corrupted session (ETX-260915-0653, ETX-260916-0044, ETX-260916-0374). No impact to other testing sessions or alternate operational Scan RDI instruments (E002017, E002225). Additional vials requested / retesting initiated. Instrument E001230 remains under observation and repair under WO-260470. |
| Impact Scoring Matrix (Quality Unit): | Frequency: 2 ¦ Severity: 2 ¦ Magnitude: 2 --> Total Risk Score: 8 Impact Level: [X] Major (3 to 17) [ ] Low [ ] Critical |

| Prepared By (Print): Qiyue Chen, Ph.D. ¦ Project Microbiologist Date: 18-Sep-2026 Signature: __________________________________ | Reviewed By (Print): Robin Seymour ¦ Microbiology Supervisor Date: ____________ Signature: __________________________________ |
| --- | --- |