Table 1: Information for ETX-261002-0288 under investigation

| Processing Analyst | Aliquoting Analyst | Sample ID | Related Microbial ID | Media with microbial growth | Microbial ID |
| --- | --- | --- | --- | --- | --- |
| Alex Saravia | Cuong Du | ETX-261002-0288 | ETX-261012-0654 | 1 x 300mL FTM | Bacillus paralicheniformis |

Table 2: Environmental Monitoring from Processing Performed on 05Oct26

| Environmental Monitoring (EM) Sampling Site | Frequency | Date (DDMMMYY) | Analyst | Day /Week(s) | Observation | EM Plate ETX Number | Microbial ID | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM |  |  |  |  |  |  |  |  |
| Personnel (Left Touch and Right Touch) | Daily | 29Sep26 | ES | Date Before Testing | No Growth | Not Applicable | None | None |
| Personnel (Left Touch and Right Touch) | Daily | 05Oct26 | ES | Date of Testing | No Growth | Not Applicable | None | None |
| Personnel (Left Touch and Right Touch) | Daily | 06Oct26 | ES | Date After Testing | No Growth | Not Applicable | None | None |
| Biological Safety Cabinet EM for E001316 |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 (4 locations) | Daily | 01Oct26 | QYC | Date Before Testing | No Growth | Not Applicable | None | None |
| Surface Sampling of ISO 5 (4 locations) | Daily | 05Oct26 | ES | Date of Testing | No Growth | Not Applicable | None | None |
| Surface Sampling of ISO 5 (4 locations) | Daily | 06Oct26 | ES | Date After Testing | No Growth | Not Applicable | None | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 01Oct26 | QYC | Date Before Testing | No Growth | Not Applicable | None | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 05Oct26 | ES | Date of Testing | No Growth | Not Applicable | None | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 06Oct26 | ES | Date After Testing | No Growth | Not Applicable | None | None |
| Weekly Active Air Sampling of Cleanroom 114, 114A, 114B (Suite 114) with Processing BSCs |  |  |  |  |  |  |  |  |
| Active Air Sampling of Cleanrooms | Weekly | 29Sep26 | SMO | Week of Testing | 3 CFUs | ETX-261006-0660 | Pending | None |
| Weekly Surface Sampling of Cleanroom 114, 114A, 114B (Suite 114) with Processing BSCs |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms | Weekly | 29Sep26 | SMO | Week of Testing | No Growth | Not Applicable | None | None |

Table 3: Environmental Monitoring from Aliquoting Performed on 12Oct26

| Environmental Monitoring (EM) Sampling Site | Frequency | Date (DDMMMYY) | Analyst | Day /Week(s) | Observation | EM Plate ETX Number | Microbial ID | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM |  |  |  |  |  |  |  |  |
| Personnel (Left Touch and Right Touch) | Daily | 09Oct26 | CCD | Date Before Testing | No Growth | Not Applicable | None | None |
| Personnel (Left Touch and Right Touch) | Daily | 12Oct26 | CCD | Date of Testing | No Growth | Not Applicable | None | None |
| Personnel (Left Touch and Right Touch) | Daily | 13Oct26 | CCD | Date After Testing | No Growth | Not Applicable | None | None |
| Biological Safety Cabinet EM for E001798 |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 (4 locations) | Daily | 09Oct26 | CCD | Date Before Testing | No Growth | Not Applicable | None | None |
| Surface Sampling of ISO 5 (4 locations) | Daily | 12Oct26 | CCD | Date of Testing | No Growth | Not Applicable | None | None |
| Surface Sampling of ISO 5 (4 locations) | Daily | 13Oct26 | CCD | Date After Testing | No Growth | Not Applicable | None | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 09Oct26 | CCD | Date Before Testing | No Growth | Not Applicable | None | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 12Oct26 | CCD | Date of Testing | No Growth | Not Applicable | None | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 13Oct26 | CCD | Date After Testing | No Growth | Not Applicable | None | None |
| Weekly Active Air Sampling of Cleanroom 114, 114A, 114B (Suite 114) with Processing BSCs |  |  |  |  |  |  |  |  |
| Active Air Sampling of Cleanrooms | Weekly | 08Oct26 | SMO | Week of Testing | No Growth | Not Applicable | None | None |
| Weekly Surface Sampling of Cleanroom 114, 114A, 114B (Suite 114) with Processing BSCs |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms | Weekly | 08Oct26 | SMO | Week of Testing | No Growth | Not Applicable | None | None |