Phase I

| # Section A – Test Details |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| # Initiator: | # Clea S. Garza (Written by: Qiyue Chen) | # Test Date: |  | # 11-May-2026 |  |  |  | # Date Initiated: |  | 18-May-2026 |  |
| # Date of Incident: | 18-May-2026 | # Test Name: |  | Environmental Monitoring |  |  |  | # Sample ID: |  | # ETX-260518-0273 |  |
| # Name of Analyst who Performed the Test: | # Clea S. Garza (Surface Sampling (Changeover) Analyst) Maraya Chukwumerije and Sophia Santamaria (Surface Sampling (Changeover) Plate Reader) | # Sample / Active Name |  | # ScanC/O CGS E001309 S1 11MAY2026 |  |  |  | # Lot #: |  | # ScanC/O CGS E001309 S1 11MAY2026 |  |
|  |  | # Dosage Form |  | # Plate |  |  |  |  |  |  |  |
| # Description of # Incident: | # The CFU count for the environmental monitoring plate exceeded the action level. |  |  |  | # Client Care notified of OOS # Sterility and Endotoxin test related OOS; Client must be notified within 24 hours of OOS |  |  |  |  |  |  |
|  | SOP / Test Method #: MICRO-SOP-2 | Effective Date: 23-Jul-2026 |  |  |  |  | SOP / Test Method Rev: 16 |  | # Limits/Specification: # <1 event |  |  |
| # Manager notified | # Manager/designee name: # Kathan Parikh | Date: 18-May-2026 |  |  |  |  | # Quality Assurance Department Notified: |  | # Yes No |  |  |
| # Section B – Investigation and Root Cause Analysis |  |  |  |  |  |  |  |  |  |  |  |
| # Personnel Discrepancy |  |  |  |  |  | # Comments |  |  |  |  |  |
| # Analyst interviewed? |  | # Yes No N/A |  |  |  | # Yes, analysts Clea S. Garza and Maraya Chukwumerije and Sophia Santamaria were comprehensively interviewed. |  |  |  |  |  |
| # Correct samples analyzed? |  | # Yes No N/A |  |  |  | # Yes, as per MICRO-SOP-2 |  |  |  |  |  |
| # Correct SOP / Test Method used/followed? |  | # Yes No N/A |  |  |  | Yes, as per MICRO-SOP-2 |  |  |  |  |  |
| # Correct sample collection/technique? |  | # Yes No N/A |  |  |  | Yes, as per MICRO-SOP-2 |  |  |  |  |  |
| # Are the Test Result Records Attached & Calculations Verified and Correct? |  | # Yes No N/A |  |  |  | # Yes, Information is available on Eagletrax under ETX-260518-0273 |  |  |  |  |  |
| # Is the analyst qualified to perform the test? |  | # Yes No N/A |  |  |  | Yes, the analysts are trained and qualified by quality to perform the test |  |  |  |  |  |
| # Is formula worksheet provided and verified? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  |  |
| # Does the processing method/ method set conform to the required specification? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  |  |
| # Was the sample stored appropriately? |  | # Yes No N/A |  |  |  | Yes, as per MICRO-SOP-2 |  |  |  |  |  |
| # Material Discrepancy |  |  |  |  |  | # Lot Number |  |  |  |  | # Expiry Date |
| # Is the integrity of the sample(s) questionable? |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |
| # Are the other test results from this run under investigation? |  | # Yes No N/A (See Phase I summary) |  |  |  |  |  |  |  |  |  |
| # Appropriate /correct glassware used? |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |
| # Are reagents stored in proper condition and not expired? |  | # Yes No N/A |  |  |  | Contact Plate: 1011543730 |  |  |  |  | Contact Plate: 29SEP2026 |
| # Standard Used |  |  |  |  |  |  |  |  |  |  |  |
| # Solvent Used |  | # Not Applicable |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Is the standard the source of the root cause? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Has the analyst been instructed to keep the original test solution for possible retesting? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Equipment Discrepancy |  |  |  |  |  | # Equipment ID |  |  |  |  | # Calibration Due Date |
| # Was the instrument set up correctly? |  | # Yes No N/A |  |  |  | Please see below |  |  |  |  | Please see below |
| # Is the clean room facility certified? |  | Yes No N/A |  |  |  | CR117 (E001739) |  |  |  |  | Dec 2026 |
| # System Performance Check (if applicable)? |  | # Yes No N/A |  |  |  | N/A |  |  |  |  | N/A |
| # System Suitability RSD < 2% (PDA)? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # System Suitability RSD < 10% (QDA,CAD)? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Do the environmental monitoring results conform to required standards? |  | Yes No N/A |  |  |  | # See Phase I Summary |  |  |  |  | # See Phase I Summary |
| # Do the controls (i.e., positive, negative, etc.) conform to the required standards? |  | Yes No N/A |  |  |  | # See Phase I Summary |  |  |  |  | # See Phase I Summary |
| # Is equipment/instrument functional and calibrated/qualified/certified/PM? |  | # Yes No N/A |  |  |  | # Incubator E001034 # (Sensor E001501) # Incubator E001031 (Sensor E001505) |  |  |  |  | # Aug26 # Feb26 # Aug26 # Feb26 |
| # Are chromatogram/spectra/raw data correct? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Section C – Phase I Summary |  |  |  |  |  |  |  |  |  |  |  |
| # Has root cause been identified as a possible laboratory error? # Note: Root cause must be conclusive, verifiable, and well-documented. # Yes # If retesting is required, complete Phase II of the OOS (initiate Form 3.100.019.F02). # No, the OOS result is valid. |  |  |  |  |  |  |  |  |  |  |  |
| Phase I Summary: # Environmental Monitoring Summary: Personnel monitoring plates for Clea S. Garza, including left- and right-touch plates, showed no microbial growth on the date before testing (08 May 2026), the date of testing (11 May 2026), and the date after testing (12 May 2026). For ISO 5 BSC E001309, daily surface sampling of four locations showed no microbial growth on the date before testing (08 May 2026). On the date of testing (11 May 2026), 1 CFU was recovered from S1 associated with Clea S. Garza. The recovery was documented under ETX-260518-0273; microbial identification indicated colony-like artifact (White, shiny, hardened, smooth area embedded on agar). Surface sampling performed on the date after testing (12 May 2026) showed no microbial growth. Settling sampling of ISO 5 BSC E001309, including two locations, showed no microbial growth on the date before testing (08 May 2026), the date of testing (11 May 2026), and the date after testing (12 May 2026). Weekly active-air monitoring of Suite 117 conducted during the week before testing and the week of testing showed no microbial growth. Weekly surface monitoring of the anteroom and cleanroom areas associated with Suite 117 showed no microbial growth during the week before testing or the week of testing. During the interview, the analyst indicated that no obvious abnormalities or deviations occurred during the testing process. All materials were disinfected before testing, and the relevant cleanroom and ISO 5 BSC were cleaned and prepared before testing in accordance with MICRO-SOP-2 and MICRO-SOP-9. Monthly cleaning and disinfection of the cleanroom suite, including the ISO 8 anteroom (117), ISO 7 buffer room (117A), ISO 7 cleanroom (117B), and the ISO 5 biosafety cabinets located within Room 117B, were performed on 26 April 2026 by Analyst - Tamiru Kotisso and Cuong Du - in accordance with MICRO-SOP-9, Cleaning and Disinfecting Procedure for Microbiology. All H2O2 indicators passed, confirming the successful completion and effectiveness of the monthly cleaning and disinfection activities within Rooms 117, 117A, and 117B. Additionally, routine cleaning and disinfection were performed before and after the testing activity in accordance with MICRO-SOP-9. It is also important to note that no samples processed by Clea S. Garza in ISO 5 BSC E001309 on the date of testing (11 May 2026) failed SCAN RDI testing. Based on the available evidence, one colony-like artifact was observed on ScanC/O CGS E001309 S1 11MAY2026 in ISO 5 BSC E001309 on the date of testing (11 May 2026). However, the observation could not be confirmed as a viable microbial CFU because no growth was obtained following inoculation onto fresh media. Therefore, the observation may represent a nonviable or non-microbial artifact, including a potential artifact associated with agar preparation or the agar-pouring process. If the observed artifact had represented a viable CFU, it would appear to be an isolated event. This assessment is supported by the absence of microbial recovery from Clea S. Garza's personnel-monitoring plates on the date before testing (08 May 2026), the date of testing (11 May 2026), and the date after testing (12 May 2026); the absence of growth from BSC E001309 surface samples collected on the date before testing and the date after testing; and the absence of growth from ISO 5 settling plates throughout the bracketing period. The negative ISO 5 settling, personnel, and bracketing surface monitoring results demonstrate that the BSC E001309 critical environment remained in a state of control. The available data do not support migration, persistence, or recurrence of contamination within ISO 5 BSC E001309. Collectively, the evidence supports that the recovery was an isolated, transient, and non-recurring event, while established cleaning, disinfection, and aseptic controls remained effective. Accordingly, no systemic environmental control deficiencies were identified, and no additional corrective or preventive actions are warranted at this time beyond continued routine environmental monitoring and adherence to approved cleaning, disinfection, and aseptic procedures. # Accordingly, no systemic environmental control deficiencies were identified, and no additional corrective or preventive actions are warranted at this time beyond continued routine environmental monitoring and adherence to approved cleaning, disinfection, and aseptic procedures. |  |  |  |  |  |  |  |  |  |  |  |
| # Reviewer Comments: |  |  |  |  |  |  |  |  |  |  |  |
| # Acknowledgement and Closure |  |  |  |  |  |  |  |  |  |  |  |
| # Can investigation be closed? # Yes # Original result is valid. # No, initiate Form 3.100.019.F02. |  |  |  |  |  |  |  |  |  |  |  |
| # Prepared by (Print): |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |
| # Lab Manager/designee (Print): |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |
| # Approved by QA Representative (Print): # Closed: YES NO |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |

# Table 1: Information for ETX-260518-0273 under investigation

# Table 2: Environmental Monitoring from Testing Performed on 11-May-2026

# Table 3: In Trend of Past OOS Results for Eagle Analytical Internal EM for analyte "ScanC/O CGS E001309 S1 11MAY2026"

Table 1: Read Dates & Incubation Observation

Table 2: Environmental Monitoring for Analyst & Cleanroom

| ETX Submission ID | Set-up Analyst & Date | Sampling Site & Location | Plate Reading Analyst (≥ 48H) | CFUs Observed after 48 Hour Incubation at 30–35°C (E001031) | Plate Reading Analyst (NLT 5 days) | CFUs Observed after NLT 5-day Incubation at 20–25°C (E001034) | Microbial Identification |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ETX-260518-0273 | CGS 11-May-2026 | Surface Sampling (Changeover) Plate (BSC E001309) | MC | No microbial growth was observed | SAS | 1 CFU on ScanC/O CGS E001309 S1 11MAY2026 | colony-like artifact (White, shiny, hardened, smooth area embedded on agar) |

| Environmental Monitoring (EM) Sampling Site | Frequency | Date (DDMMMYYYY) | Analyst (Initials) | Day /Week(s) | Observation | Environmental Monitoring Plate ETX ID | Microbial ID | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM Bracketing |  |  |  |  |  |  |  |  |
| Personal (Left Touch and Right Touch) | Daily | 08 May 2026 | CGS | Date Before Testing | No growth | N/A | N/A | None |
| Personal (Left Touch and Right Touch) | Daily | 11-May-2026 | CGS | Date of Testing | No growth | N/A | N/A | None |
| Personal (Left Touch and Right Touch) | Daily | 12 May 2026 | CGS | Date After Testing | No growth | N/A | N/A | None |
| Biological Safety Cabinet EM Bracketing Biological Safety Cabinet (BSC) |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 (4 locations) | Daily | 08 May 2026 | CGS | Date Before Testing | No growth | N/A | N/A | None |
| Surface Sampling of ISO 5 (4 locations) | Daily | 11-May-2026 | CGS | Date of Testing | 1 CFU | ETX-260518-0273 | colony-like artifact (White, shiny, hardened, smooth area embedded on agar) | None |
| Surface Sampling of ISO 5 (4 locations) | Daily | 12 May 2026 | CGS | Date After Testing | No growth | N/A | N/A | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 08 May 2026 | CGS | Date Before Testing | No growth | N/A | N/A | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 11-May-2026 | CGS | Date of Testing | No growth | N/A | N/A | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 12 May 2026 | CGS | Date After Testing | No growth | N/A | N/A | None |
| Weekly Active Air Sampling Bracketing |  |  |  |  |  |  |  |  |
| Active Air Sampling of Cleanrooms | Weekly | 08 May 2026 | Rey Estrada | Week (On or After Testing Date) | No growth | N/A | N/A | None |
| Surface Sampling of Anteroom and Cleanroom Bracketing |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms | Weekly | 08 May 2026 | Rey Estrada | Week (On or After Testing Date) | No growth | N/A | N/A | None |