EAGLE ANALYTICAL SERVICES
CHANGE CONTROL REQUEST FORM (CORP-FORM-4 / 3.100.004.F01 Rev 10)
CR Number: CR-26449 | Coupled with: OOS-261814 & NCR N26414

―――――――――――――――――――――――――――――――――――――――――――――――――――――――――――――――――

1. Request Originator & Reason for Change

| Name: Qiyue Chen | Department: Microbiology |
| --- | --- |
| Date Requested: 17-Sep-2026 | Change Control #: CR-26449 |
| Type of Change: Other: EagleTrax Test Result Correction (Pass to Fail) | Affected System: EagleTrax LIMS |
| Reason for Change: OOS #OOS-261814 & NCR #N26414 (Sample result transposition correction) | Coupled With: OOS-261814 (Phase I OOS Report) |

2. Change Description

| ETX-260805-0189 Client: GoGoMeds Select (E10747) Lot #: 18468 Analyte: Tirzepatide/B12 25/1mg/mL Injectable Test: Scan RDI Sterility Coupled Quality Records: • OOS-261814 • NCR N26414 | On 07-Aug-2026, sample ETX-260805-0189 (GoGoMeds Select (E10747), Lot 18468) failed Scan RDI sterility testing with a positive result of 4 CFUs (rod morphology). Due to an unintentional clerical transposition during manual data entry, this failing result was inadvertently approved as 'Passing/Complete' in EagleTrax, while passing sample ETX-260804-0101 was placed on hold. This nonconformance was immediately identified and documented under NCR N26414. Eagle Client Care and Laboratory Management promptly contacted the client on 07-Aug-2026 to notify them of the positive test result and confirm immediate quarantine of Lot 18468, preventing unintended product release. Subsequent Phase I laboratory investigation OOS-261814 confirmed that the positive sterility result was valid and not caused by laboratory contamination or analytical error. Retraining on MICRO-SOP-12 and CORP-SOP-15 was completed under NCR N26414 in ZenQMS. Required Change: Within Eagle Trax Submission Information, under the Test Results field for Submission ID ETX-260805-0189, update the entry as follows: • Change 'Pass' to 'Fail' (Positive: 4 CFUs, rod-shaped morphology). • Cross-reference both closed investigation OOS-261814 and closed NCR N26414 within ETX test notes. This change ensures the submission record accurately reflects the validated failing test result and aligns EagleTrax with finalized quality records. |
| --- | --- |

3. Impact of Change

| Frequency | Severity | Magnitude | Impact Grading Score & Level |
| --- | --- | --- | --- |
| <2 instances = 1 | Changes with retroactive implication = 3 | <2 Documents/Processes = 1 | Score: 1 × 3 × 1 = 3 Level: Major (Score 3 to 17) |

The transposition occurred during manual data entry on 07-Aug-2026. Client Care immediately notified the client on the same day to quarantine Lot 18468, preventing release of affected product. Investigation OOS-261814 confirmed test validity. This CR authorizes correcting EagleTrax records from Pass to Fail. Retraining was completed under closed NCR N26414. Overall product quality risk is fully controlled.

4. Reviewed & Approved

| Department | Name | Title | Date |
| --- | --- | --- | --- |
| Microbiology | Robin Seymour | Microbiology Supervisor (Sterile Lab) | 17-Sep-2026 |
| Microbiology | Kathan Parikh | Associate Director of Microbiology | 17-Sep-2026 |

Quality Closure (Completed by Quality)

Training Required: N/A (Documented Read & Understood retraining on MICRO-SOP-12 and CORP-SOP-15 was previously assigned and completed under Nonconformance Report NCR N26414 in ZenQMS).

Closure Comments: Comments: Result correction in EagleTrax coupled with closed OOS-261814 and closed NCR N26414. Ready for final Quality closure in accordance with CORP-SOP-4.