Phase I

| # Section A – Test Details |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| # Initiator: | # {{ analyst_signature }} | # Test Date: |  | # {{ test_date }} |  |  |  | # Date Initiated: |  | # {{ test_date }} |  |
| # Date of Incident: | # {{ test_date }} | # Test Name: |  | # Scan RDI Sterility Test |  |  |  | # Sample ID: |  | # {{ report_header }} |  |
| # Name of Analyst who Performed the Test: | # {{ smart_personnel_block }} | # Sample / Active Name |  | # {{ sample_name }} |  |  |  | # Lot #: |  | # {{ lot_number }} |  |
|  |  | # Dosage Form |  | # {{ dosage_form}} |  |  |  |  |  |  |  |
| # Description of # Incident: | # {{ smart_incident_opening }} |  |  |  | # Client Care notified of OOS # Sterility and Endotoxin test related OOS; Client must be notified within 24 hours of OOS |  |  |  |  |  |  |
|  | SOP / Test Method #: MICRO-SOP-12 (16) ENG-SOP-4 (05) | Effective Date: 24Jul26 24Jul26 |  |  |  |  | SOP / Test Method Rev: Rev: 16 Rev: 05 |  | # Limits/Specification: # <1 event |  |  |
| # Manager notified | # Manager/designee name: # Kathan Parikh | # Date: # {{ test_date }} |  |  |  |  | # Quality Assurance Department Notified: |  | # Yes No |  |  |
| # Section B – Investigation and Root Cause Analysis |  |  |  |  |  |  |  |  |  |  |  |
| # Personnel Discrepancy |  |  |  |  |  | # Comments |  |  |  |  |  |
| # Analyst interviewed? |  | # Yes No N/A |  |  |  | # {{ smart_comment_interview }} |  |  |  |  |  |
| # Correct samples analyzed? |  | # Yes No N/A |  |  |  | # {{ smart_comment_samples }} |  |  |  |  |  |
| # Correct SOP / Test Method used/followed? |  | # Yes No N/A |  |  |  | Yes, as per MICRO-SOP-12, ENG-SOP-4 |  |  |  |  |  |
| # Correct sample collection/technique? |  | # Yes No N/A |  |  |  | Yes, as per MICRO-SOP-12, ENG-SOP-4 |  |  |  |  |  |
| # Are the Test Result Records Attached & Calculations Verified and Correct? |  | # Yes No N/A |  |  |  | # {{ smart_comment_records }} |  |  |  |  |  |
| # Is the analyst qualified to perform the test? |  | # Yes No N/A |  |  |  | # Yes, all analysts are trained and qualified by quality to perform the test. |  |  |  |  |  |
| # Is formula worksheet provided and verified? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  |  |
| # Does the processing method/ method set conform to the required specification? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  |  |
| # Was the sample stored appropriately? |  | # Yes No N/A |  |  |  | # {{ smart_comment_storage }} |  |  |  |  |  |
| # Material Discrepancy |  |  |  |  |  | # Lot Number |  |  |  |  | # Expiry Date |
| # Is the integrity of the sample(s) questionable? |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |
| # Are the other test results from this run under investigation? |  | # Yes No N/A (See Phase I summary) |  |  |  |  |  |  |  |  |  |
| # Appropriate /correct glassware used? |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |
| # Are reagents stored in proper condition and not expired? |  | # Yes No N/A |  |  |  | # Scan Consumables: # See the attached data packet # Environmental Plates: # TSA and Surface Plate: see attached environmental logs # Monthly Cleaning: # H2O2 strips and IPA: See attached monthly cleaning logs |  |  |  |  | # Scan Consumables: # See attached data packet # Environmental Plates: # TSA and Surface Plate: see attached environmental logs # Monthly Cleaning: # H2O2 strips and IPA: See attached monthly cleaning logs |
| # Standard Used |  | # {{ control_positive }} |  |  |  | # {{ control_lot }} |  |  |  |  | # {{ control_data }} |
| # Solvent Used |  | # Not Applicable |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Is the standard the source of the root cause? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Has the analyst been instructed to keep the original test solution for possible retesting? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Equipment Discrepancy |  |  |  |  |  | # Equipment ID |  |  |  |  | # Calibration Due Date |
| # Was the instrument set up correctly? |  | # Yes No N/A |  |  |  | # E00{{ scan_id }} |  |  |  |  | # Oct26 |
| # Is the clean room facility certified? |  | Yes No N/A |  |  |  | # E00{{ cr_id }} (CR{{ cr_suit }}) |  |  |  |  | # Dec25 |
| # System Performance Check (if applicable)? |  | # Yes No N/A |  |  |  | # E00{{ scan_id }} |  |  |  |  | # Oct26 |
| # System Suitability RSD < 2% (PDA)? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # System Suitability RSD < 10% (QDA,CAD)? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Do the environmental monitoring results conform to required standards? |  | Yes No N/A |  |  |  | # See Phase I Summary |  |  |  |  | # See Phase I Summary |
| # Do the controls (i.e., positive, negative, etc.) conform to the required standards? |  | Yes No N/A |  |  |  | # See Phase I Summary |  |  |  |  | # See Phase I Summary |
| # Is equipment/instrument functional and calibrated/qualified/certified/PM? |  | # Yes No N/A |  |  |  | # Incubator E001034 # (Sensor E001501) # Incubator E001031 (Sensor E001505) |  |  |  |  | # Aug26 # Feb26 # Aug26 # Feb26 |
| # Are chromatogram/spectra/raw data correct? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Section C – Phase I Summary |  |  |  |  |  |  |  |  |  |  |  |
| # Has root cause been identified as a possible laboratory error? # Note: Root cause must be conclusive, verifiable, and well-documented. # Yes # If retesting is required, complete Phase II of the OOS (initiate Form 3.100.019.F02). # No, the OOS result is valid. |  |  |  |  |  |  |  |  |  |  |  |
| Phase I Summary: # {{ smart_phase1_summary }} # {{ smart_phase1_continued }} |  |  |  |  |  |  |  |  |  |  |  |
| # Reviewer Comments: |  |  |  |  |  |  |  |  |  |  |  |
| # Acknowledgement and Closure |  |  |  |  |  |  |  |  |  |  |  |
| # Can investigation be closed? # Yes # Original result is valid. # No, initiate Form 3.100.019.F02. |  |  |  |  |  |  |  |  |  |  |  |
| # Prepared by (Print): |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |
| # Lab Manager/designee (Print): |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |
| # Approved by QA Representative (Print): # Closed: YES NO |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |

# Table 1: Environmental Monitoring from {{ sample_id }}.

| EM Sampling Site | Frequency | Date DDMMM YY | Analyst | Day /Week(s) | Observation | EM Plate ETX Number | Microbial ID | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM Bracketing for {{ analyst_name }} |  |  |  |  |  |  |  |  |
| Personal (Left Touch and Right Touch) | Daily | {{ test_date }} | {{ analyst_initial }} | Date of Testing | {{ obs_pers_dur }} | {{ etx_pers_dur }} | {{ id_pers_dur }} | None |
| Biological Safety Cabinet EM Bracketing for BSC E00{{ bsc_id }} |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 BSC E00{{ bsc_id }} (4 locations) | Daily | {{ test_date }} | {{ analyst_initial }} | Date of Testing | {{ obs_surf_dur }} | {{ etx_surf_dur }} | {{ id_surf_dur }} | None |
| Settling Sampling of ISO 5 BSC E00{{ bsc_id }} | Daily | {{ test_date }} | {{ analyst_initial }} | Date of Testing | {{ obs_sett_dur }} | {{ etx_sett_dur }} | {{ id_sett_dur }} | None |
| Weekly Active Air Sampling Bracketing for CR{{ cr_id }} |  |  |  |  |  |  |  |  |
| Active Air Sampling of Cleanrooms | Weekly | {{ date_of_weekly }} | {{ weekly_initial }} | Week of Testing | {{ obs_air_wk_of }} | {{ etx_air_wk_of }} | {{ id_air_wk_of }} | None |
| Surface Sampling of Anteroom and Cleanroom Bracketing for CR{{ cr_id }} |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms | Weekly | {{ date_of_weekly }} | {{ weekly_initial }} | Week of Testing | {{ obs_room_wk_of }} | {{ etx_room_wk_of }} | {{ id_room_wk_of }} | None |

*Description of CFU# and location where growth was observed.

N/A if Not Applicable