Table 1: Read Dates & Incubation Observation

| ETX Submission ID | Set-up Analyst & Date | Sampling Site & Location | Plate Reading Analyst (≥ 48H) | CFUs Observed after 48 Hour Incubation at 30–35°C (E001031) | Plate Reading Analyst (NLT 5 days) | CFUs Observed after NLT 5-day Incubation at 20–25°C (E001034) | Microbial Identification |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ETX-260518-0273 | CGS 11MAY 2026 | Surface sampling plate | MC | 0 | SAS | 1 CFU on surface #1 sampling plate | N/A |

Table 2: Environmental Monitoring Plates for Analyst and Cleanroom Bracketing

| Environmental Monitoring (EM) Sampling Site | Frequency | Date (DDMM YYYY) | Analyst (Initials) | Day /Week(s) | Observation* | Environmental Monitoring Plate ETX ID | Microbial ID | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM Bracketing |  |  |  |  |  |  |  |  |
| Personal (Left Touch and Right Touch) | Daily | 08MAY 2026 | CGS | Date Before Testing | No growth | N/A | N/A | None |
| Personal (Left Touch and Right Touch) | Daily | 11MAY 2026 | CGS | Date of Testing | No growth | N/A | N/A | None |
| Personal (Left Touch and Right Touch) | Daily | 12MAY 2026 | CGS | Date After Testing | No growth | N/A | N/A | None |
| Biological Safety Cabinet EM Bracketing Biological Safety Cabinet (BSC) 1309 |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 (4 locations) | Daily | 09MAY 2026 | VV | Date Before Testing | No growth |  |  | None |
| Surface Sampling of ISO 5 (4 locations) | Daily | 11MAY 2026 | CGS, DT, GA, VV | Date of Testing | 1 CFU on S1 for CGS | ETX-260518-0273 | N/A White, shiny, hardened, smooth area embeded on agar | None |
| Surface Sampling of ISO 5 (4 locations) | Daily | 12MAY 2026 | PG, DT, HS, ELB, GA, VV | Date After Testing | No growth | N/A | N/A | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 09MAY 2026 | VV | Date Before Testing | No growth | N/A | N/A | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 11MAY 2026 | CGS, DT, GA, VV | Date of Testing | No growth | N/A | N/A | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 12MAY 2026 | PG, DT, HS, ELB, GA, VV | Date After Testing | No growth | N/A | N/A | None |
| Weekly Active Air Sampling Bracketing 117 |  |  |  |  |  |  |  |  |
| Active Air Sampling of Cleanrooms | Weekly | 29APR 2026 | SMO | Week (On of Testing Date) | No growth | N/A | N/A | None |
| Active Air Sampling of Cleanrooms | Weekly | 07MAY 2026 | SMO | Week (On or After Testing Date) | 2 CFU for 117 Air | ETX-260518-0263 | Gram (+) cocci Gram (+) pleomorphic rods | None |
| Surface Sampling of Anteroom and Cleanroom Bracketing 117 |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms | Weekly | 29APR 2026 | SMO | Week (On of Testing Date) | No growth | N/A | N/A | None |
| Surface Sampling of Cleanrooms | Weekly | 07MAY 2026 | SMO | Week (On or After Testing Date) | No growth | N/A | N/A | None |