# Table 1: Information for ETX-260828-0527 under investigation

| Processing Analyst | Aliquoting Analyst | Sample ID | Related Microbial ID | Media with microbial growth | Microbial ID |
| --- | --- | --- | --- | --- | --- |
| Gabrielle Surber | Cuong Du | ETX-260828-0527 | ETX-260908-0593 | 1 x 300mL FTM | No growth was obtained |

# Table 2: Environmental Monitoring from Processing Performed on 01Sep26

| Environmental Monitoring (EM) Sampling Site | Frequency |  | Date (DDMMMYY) | Analyst (Initials) |  | Day /Week(s) | Observation* |  | Environmental Monitoring Plate ETX ID | Microbial ID | Notes |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM Bracketing |  |  |  |  |  |  |  |  |  |  |  |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 31Aug26 | GS |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 01Sep26 | GS |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 02Sep26 | GS |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Biological Safety Cabinet EM Bracketing Biological Safety Cabinet (BSC) |  |  |  |  |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 31Aug26 | GS |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 01Sep26 | GS |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 02Sep26 | GS |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 31Aug26 | GS |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 01Sep26 | GS |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 02Sep26 | GS |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Weekly Active Air Sampling Bracketing |  |  |  |  |  |  |  |  |  |  |  |  |
| Active Air Sampling of Cleanrooms |  | Weekly | 04Sep26 | ISS | Week on Testing Date |  |  | 1 CFU (ISO 8 114) | ETX-260914-0487 | Corynebacterium ureicelerivorans Mycobacterium grossiae |  | None |
| Surface Sampling of Anteroom and Cleanroom Bracketing |  |  |  |  |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms |  | Weekly | 04Sep26 | ISS | Week on Testing Date |  |  | No growth | N/A | N/A |  | None |

# Table 3: Environmental Monitoring from Aliquoting Performed on 08Sep26

| Environmental Monitoring (EM) Sampling Site | Frequency |  | Date (DDMMMYY) | Analyst (Initials) |  | Day /Week(s) | Observation* |  | Environmental Monitoring Plate ETX ID | Microbial ID | Notes |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM Bracketing |  |  |  |  |  |  |  |  |  |  |  |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 03Sep26 | CCD |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 08Sep26 | CCD |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 15Sep26 | CCD |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Biological Safety Cabinet EM Bracketing Biological Safety Cabinet (BSC) |  |  |  |  |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 07Sep26 | ALA |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 08Sep26 | CCD |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 09Sep26 | ALA |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 07Sep26 | ALA |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 08Sep26 | CCD |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 09Sep26 | ALA |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Weekly Active Air Sampling Bracketing |  |  |  |  |  |  |  |  |  |  |  |  |
| Active Air Sampling of Cleanrooms |  | Weekly | 10Sep26 | SMO | Week on Testing Date |  |  | 1 CFU (ISO 8 114) | ETX-260921-0520 | Micrococcus luteus |  | None |
| Surface Sampling of Anteroom and Cleanroom Bracketing |  |  |  |  |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms |  | Weekly | 10Sep26 | SMO | Week on Testing Date |  |  | No growth | N/A | N/A |  | None |