Phase I

| # Section A – Test Details |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| # Initiator: | # Alex Saravia, Abayomi Odugbesi # (Written by: Qiyue Chen) | # Test Date: |  | # 04Sep26, 15Sep26 |  |  |  | # Date Initiated: |  | # 10Sep26, 21Sep26 |  |
| # Date of Incident: | # 10Sep26, 21Sep26 | # Test Name: |  | # USP <71> Sterility Test |  |  |  | # Sample ID: |  | # ETX-260902-0505, ETX-260914-0470 # Solyn LLC (E75000) |  |
| # Name of Analyst who Performed the Test: | # Prepper: # Alex Saravia, Andrew Carrillo (ES, AC) # Processor: Alex Saravia, Abayomi Odugbesi (ES, AOD) # Subculture Processor: () # Reading Analyst: # Elysse Nioupin, Andrew Carrillo (EN, AC) | # Sample / Active Name |  | # GLP3R/Cagrilinitide |  |  |  | # Lot #: |  | # 2608-216 |  |
|  |  | # Dosage Form |  | # Liquid |  |  |  |  |  |  |  |
| # Description of # Incident: | # On 10Sep26, 21Sep26, sample ETX-260902-0505, ETX-260914-0470 was found positive for viable microorganisms after USP <71> sterility testing. |  |  |  | # Client Care notified of OOS # Sterility and Endotoxin test related OOS; Client must be notified within 24 hours of OOS |  |  |  |  |  |  |
|  | # SOP / Test Method #: # MICRO-SOP-5 (19) | # Effective Date: # 24Jul26 |  |  |  |  | # SOP / Test Method Rev: # Rev: 19 |  | # Limits/Specification: # Pass/Fail |  |  |
| # Manager notified | # Manager/designee name: # Kathan Parikh | # Date: # 10Sep26, 21Sep26 |  |  |  |  | # Quality Assurance Department Notified: |  | # Yes No |  |  |
| # Section B – Investigation and Root Cause Analysis |  |  |  |  |  |  |  |  |  |  |  |
| # Personnel Discrepancy |  |  |  |  |  | # Comments |  |  |  |  |  |
| # Analyst interviewed? |  | # Yes No N/A |  |  |  | # Yes, analysts Alex Saravia, Andrew Carrillo, Alex Saravia, Abayomi Odugbesi, , and Elysse Nioupin, Andrew Carrillo were interviewed comprehensively. |  |  |  |  |  |
| # Correct samples analyzed? |  | # Yes No N/A |  |  |  | # Yes, sample ID: ETX-260902-0505, ETX-260914-0470 |  |  |  |  |  |
| # Correct SOP / Test Method used/followed? |  | # Yes No N/A |  |  |  | # Yes, as per MICRO-SOP-5 |  |  |  |  |  |
| # Correct sample collection/technique? |  | # Yes No N/A |  |  |  | # Yes, as per MICRO-SOP-5 |  |  |  |  |  |
| # Are the Test Result Records Attached & Calculations Verified and Correct? |  | # Yes No N/A |  |  |  | # Yes, Information is available in EagleTrax under ETX-260902-0505, ETX-260914-0470 |  |  |  |  |  |
| # Is the analyst qualified to perform the test? |  | # Yes No N/A |  |  |  | # Yes, all analysts are trained and qualified by quality to perform the test. |  |  |  |  |  |
| # Is formula worksheet provided and verified? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  |  |
| # Does the processing method/ method set conform to the required specification? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  |  |
| # Was the sample stored appropriately? |  | # Yes No N/A |  |  |  | # Yes, the sample was stored as per client's instructions. Information is available in EagleTrax Sample Location History under ETX-260902-0505, ETX-260914-0470 |  |  |  |  |  |
| # Material Discrepancy |  |  |  |  |  | # Lot Number |  |  |  |  | # Expiry Date |
| # Is the integrity of the sample(s) questionable? |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |
| # Are the other test results from this run under investigation? |  | # Yes No N/A (See Phase I summary) |  |  |  |  |  |  |  |  |  |
| # Appropriate /correct glassware used? |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |
| # Are reagents stored in proper condition and not expired? |  | # Yes No N/A |  |  |  | # Process Consumables: # Information is available in EagleTrax Sample Consumables section # Environmental Plates: # TSA and Surface Plate: see attached environmental logs # Monthly Cleaning: # H2O2 strips and IPA: See attached monthly cleaning logs |  |  |  |  | # Process Consumables: # Information is available in EagleTrax Sample Consumables section # Environmental Plates: # TSA and Surface Plate: see attached environmental logs # Monthly Cleaning: # H2O2 strips and IPA: See attached monthly cleaning logs |
| # Standard Used |  | # Not Applicable |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Solvent Used |  | # Not Applicable |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Is the standard the source of the root cause? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Has the analyst been instructed to keep the original test solution for possible retesting? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Equipment Discrepancy |  |  |  |  |  | # Equipment ID |  |  |  |  | # Calibration Due Date |
| # Was the instrument set up correctly? |  | # Yes No N/A |  |  |  | # ISO 5 BSC E001316, 1314 |  |  |  |  | # Mar 2026 |
| # Is the clean room facility certified? |  | Yes No N/A |  |  |  | # E001736 (CR114, 115) |  |  |  |  | # Jun 2026 # Dec 2026 |
| # System Performance Check (if applicable)? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Mar 2026 |
| # System Suitability RSD < 2% (PDA)? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # System Suitability RSD < 10% (QDA,CAD)? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Do the environmental monitoring results conform to required standards? |  | Yes No N/A |  |  |  | # See Phase I Summary |  |  |  |  | # See Phase I Summary |
| # Do the controls (i.e., positive, negative, etc.) conform to the required standards? |  | Yes No N/A |  |  |  | # See Phase I Summary |  |  |  |  | # See Phase I Summary |
| # Is equipment/instrument functional and calibrated/qualified/certified/PM? |  | # Yes No N/A |  |  |  | # Incubator E001034 # (Sensor E001501) # Incubator E001031 (Sensor E001505) # Incubator E001356 (Sensor E001450) # Incubator E001357 (Sensor E001449) |  |  |  |  | # Jun 2026 # Feb 2027 # Jun 2026 # Feb 2027 # Jun 2026 # Feb 2027 |
| # Are chromatogram/spectra/raw data correct? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Section C – Phase I Summary |  |  |  |  |  |  |  |  |  |  |  |
| # Has root cause been identified as a possible laboratory error? # Note: Root cause must be conclusive, verifiable, and well-documented. # Yes # If retesting is required, complete Phase II of the OOS (initiate Form 3.100.019.F02). # No, the OOS result is valid. |  |  |  |  |  |  |  |  |  |  |  |
| # Phase I Summary: All analysts involved in the prepping, processing, and reading of both samples – Alex Saravia, Elysse Nioupin, Andrew Carrillo, and Abayomi Odugbesi – were interviewed comprehensively. Their answers are recorded throughout this document. Upon arrival, both sample submissions (ETX-260902-0505 and ETX-260914-0470) were stored in accordance with the Client’s instructions. Analysts verified the integrity of the sample containers throughout both preparation and processing stages. No leaks, cracks, or turbidity were observed prior to testing, verifying container integrity. All reagents and supplies mentioned in the material section above were stored according to suppliers’ recommendations, and their integrity was visually verified before utilization. Moreover, all culture media, rinses, and supplies possessed valid expiration dates. The functionality of all equipment was confirmed by reviewing continuous monitoring data. Initial testing was conducted in Cleanroom Suite 114 (comprising the innermost ISO 7 cleanroom 114B, middle ISO 7 buffer room 114A, and outermost ISO 8 anteroom 114). Retest processing was conducted in Cleanroom Suite 115. Positive air pressure cascades are maintained throughout both suites to ensure controlled, outward unidirectional airflow. Initial sample ETX-260902-0505 was received on 02 Sep 2026 and processed on 04 Sep 2026 by Alex Saravia in certified ISO 5 BSC E001316 (Cleanroom 114B) as per MICRO-SOP-5 (USP 71> / EP 2.6.1 Sterility Test). Prior to entry, vials were disinfected with acidified bleach allowing a 10-minute contact time at each transfer boundary. Membrane filtration was performed (3 vials filtered). Media bottles were transferred into designated incubators E001356 and E001357. On 10 Sep 2026 (Day 6 of incubation), turbidity was observed in 1 x 100mL TSB bottle by reading analyst Elysse Nioupin and confirmed by supervisor Robin Seymour. The positive bottle was submitted for Microbial Identification under ETX-260910-0290, which definitively identified the isolate as Microbacterium sp. PM5 (Gram (+) rods). Concurrent negative controls remained sterile. Secondary sample ETX-260914-0470 was prepped on 14 Sep 2026 by Andrew Carrillo and processed on 15 Sep 2026 by Abayomi Odugbesi in certified ISO 5 BSC E001314 in Cleanroom Suite 115 as per MICRO-SOP-5 via membrane filtration (4 vials / 60 gm filtered). Standard multi-barrier disinfection and strict aseptic techniques were maintained. Media bottles were incubated in incubators E001356 and E001357. On 21 Sep 2026 (Day 6 of incubation), turbidity was observed in 1 x 100mL TSB bottle by reading analyst Andrew Carrillo and confirmed by supervisor Robin Seymour. The positive bottle was submitted for Microbial Identification under ETX-260921-0498, which definitively identified the identical isolate, Microbacterium sp. PM5 (Gram (+) short rods). Concurrent negative controls remained sterile. ISO 5 BSC Environmental Monitoring Results Evaluation: After reviewing the Environmental Monitoring (EM) records for both testing dates (04 Sep 2026 and 15 Sep 2026), no microbial growth was detected on operator personnel fingertip touch plates, settling plates, or ISO 5 BSC surface contact plates for either session, nor on preceding or subsequent testing days (including 03Sep26, 04Sep26, 08Sep26 for ES in BSC E001316, and 14Sep26, 15Sep26, 16Sep26 for AOD in BSC E001314). Cleanroom Environmental Monitoring Results Evaluation: For the initial processing on 04 Sep 2026 in Suite 114, weekly active surface sampling showed no growth; weekly active air sampling in ISO 8 cleanroom 114 recovered 2 CFUs (ETX-260914-0487), identified as Corynebacterium ureicelerivorans and Mycobacterium grossiae; on 10 Sep 2026, 1 CFU was recovered (ETX-260921-0520), identified as Micrococcus luteus. These environmental isolates differed distinctly in genus and morphology from the sample isolate (Microbacterium sp. PM5). For the retest processing on 15 Sep 2026 in Cleanroom Suite 115, weekly active surface sampling showed no growth; weekly active air sampling in ISO 8 cleanroom 115 recovered 9 CFUs (ETX-260923-0402), identified as Kocuria indica, Brevibacterium sp. CS2, Corynebacterium sp, Micrococcus luteus, Paracoccus yeei, Staphylococcus hominis, and Kocuria rhizophila. 100% of these environmental isolates differed distinctly in genus and morphology from the product contamination isolate (Microbacterium sp. PM5). Both air recoveries were strictly confined to the outermost ISO 8 anterooms (ISO 8 114 and ISO 8 115). All actual testing occurred within certified ISO 5 BSCs (E001316 in Suite 114 and E001314 in Suite 115) located in ISO 7 cleanrooms separated by positive pressure cascades. Test containers were transferred in disinfected, lidded bins, and critical zone settling, surface contact, and analyst fingertip plates remained 100% sterile (0 CFU). Therefore, cleanroom air recoveries were unrelated to the sample contamination. (Refer to Table 1 for sample details, Table 2 for environmental monitoring from processing performed on 04 Sep 2026, and Table 3 for environmental monitoring from processing performed on 15 Sep 2026). The complete absence of contamination across all analyst glove touch plates, settling plates, and ISO 5 work surfaces across both processing sessions confirms that the ISO 5 critical zones remained in optimal control. Full compliance with MICRO-SOP-9 and MICRO-SOP-5 was verified. Monthly cleaning and disinfection of Cleanroom Suite 114 and Suite 115 was performed on 30 Aug 2026 as per MICRO-SOP-9, with all chemical H2O2 indicators passing. A 6-month historical review for client Solyn LLC (E75000) confirmed that analyte GLP3R/Cagrilinitide had no prior sterility failures. Batch cross-contamination reviews were conducted for all samples processed on 04 Sep 2026 (where ETX-260902-0505 was 3rd in the batch) and 15 Sep 2026. All other client samples processed during both testing sessions tested negative, ruling out cross-contamination. Importantly, the investigation evaluated why the prior rapid ScanRDI sterility test on this lot (ETX-260807-0602) yielded a passing result. In compounded parenteral products, low-level bioburden is characteristically non-uniformly distributed across individual vials. Under low-concentration non-uniform contamination, random unit sampling naturally results in some vials containing zero viable cells—as occurred with the vials sampled for ScanRDI under ETX-260807-0602—while other vials harbor low-level viable bioburden. Furthermore, no method suitability was on file for ScanRDI with this formulation. When low-level viable organisms were introduced into the 14-day USP 71> membrane filtration enrichment process, the microorganisms were successfully enriched and recovered in TSB on Day 6 in both independent testing events (ETX-260902-0505 and ETX-260914-0470), yielding the identical strain (Microbacterium sp. PM5). The successful replication across different processing analysts (Alex Saravia vs. Abayomi Odugbesi), different BSCs (E001316 vs. E001314), and different cleanroom suites (Suite 114 vs. Suite 115) conclusively refutes the preliminary hypothesis of analyst handling error during reconstitution. The passing ScanRDI result and both failing USP 71> results are scientifically consistent with non-uniform microbial distribution. Consequently, inherent product contamination in Lot 2608-216 is confirmed, and Lot 2608-216 fails USP 71> sterility requirements. |  |  |  |  |  |  |  |  |  |  |  |
| # Reviewer Comments: |  |  |  |  |  |  |  |  |  |  |  |
| # Acknowledgement and Closure |  |  |  |  |  |  |  |  |  |  |  |
| # Can investigation be closed? # Yes # Original result is valid. # No, initiate Form 3.100.019.F02. |  |  |  |  |  |  |  |  |  |  |  |
| # Prepared by (Print): |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |
| # Lab Manager/designee (Print): |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |
| # Approved by QA Representative (Print): # Closed: YES NO |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |

# Table 1: Information for ETX-260902-0505, ETX-260914-0470 under investigation

| Processing Analyst | Reading Analyst | Sample ID | Related Microbial ID | Media with microbial growth | Microbial ID |
| --- | --- | --- | --- | --- | --- |
| Alex Saravia | Elysse Nioupin | ETX-260902-0505 | ETX-260910-0290 | 1 x 100mL TSB bottle | Microbacterium sp. PM5 (Gram (+) rods) |
| Abayomi Odugbesi | Andrew Carrillo | ETX-260914-0470 | ETX-260921-0498 | 1 x 100mL TSB bottle | Microbacterium sp. PM5 (Gram (+) short rods) |

# Table 2: Environmental Monitoring from Processing Performed on 04Sep26

| Environmental Monitoring (EM) Sampling Site | Frequency |  | Date (DDMMMYY) | Analyst (Initials) |  | Day /Week(s) | Observation* |  | Environmental Monitoring Plate ETX ID | Microbial ID | Notes |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM Bracketing |  |  |  |  |  |  |  |  |  |  |  |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 03Sep26 | ES |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 04Sep26 | ES |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 08Sep26 | ES |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Biological Safety Cabinet EM Bracketing - ISO 5 BSC E001316 (Suite 114) |  |  |  |  |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 03Sep26 | ES |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 04Sep26 | ES |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 08Sep26 | ES |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 03Sep26 | ES |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 04Sep26 | ES |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 08Sep26 | ES |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Weekly Active Air Sampling Bracketing (Suite 114) |  |  |  |  |  |  |  |  |  |  |  |  |
| Active Air Sampling of Cleanrooms |  | Weekly | 04Sep26 | ISS | Week on Testing Date |  |  | 2 CFUs (ISO 8 114) | ETX-260914-0487 | Corynebacterium ureicelerivorans Mycobacterium grossiae |  | None |
| Weekly Surface Sampling of Cleanrooms Bracketing (Suite 114) |  |  |  |  |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms |  | Weekly | 04Sep26 | ISS | Week on Testing Date |  |  | No growth | N/A | N/A |  | None |

Table 3: Environmental Monitoring from Processing Performed on 15Sep26

| Environmental Monitoring (EM) Sampling Site | Frequency |  | Date (DDMMMYY) | Analyst (Initials) |  | Day /Week(s) | Observation* |  | Environmental Monitoring Plate ETX ID | Microbial ID | Notes |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM Bracketing |  |  |  |  |  |  |  |  |  |  |  |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 14Sep26 | AOD |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 15Sep26 | AOD |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Personnel (Left Touch and Right Touch) | Daily |  | 16Sep26 | AOD |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Biological Safety Cabinet EM Bracketing - ISO 5 BSC E001314 (Suite 115) |  |  |  |  |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 14Sep26 | AOD |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 15Sep26 | AOD |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Surface Sampling of ISO 5 (4 locations) | Daily |  | 16Sep26 | AOD |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 14Sep26 | AOD |  | Date Before Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 15Sep26 | AOD |  | Date of Testing | No growth |  | N/A | N/A | None |  |
| Settling Sampling of ISO 5 (2 locations) | Daily |  | 16Sep26 | AOD |  | Date After Testing | No growth |  | N/A | N/A | None |  |
| Weekly Active Air Sampling Bracketing (Suite 115) |  |  |  |  |  |  |  |  |  |  |  |  |
| Active Air Sampling of Cleanrooms |  | Weekly | 15Sep26 | SMO | Week on Testing Date |  |  | 9 CFUs (ISO 8 115) | ETX-260923-0402 | Kocuria indica Brevibacterium sp. CS2 Corynebacterium sp Micrococcus luteus Paracoccus yeei Staphylococcus hominis Kocuria rhizophila |  | None |
| Weekly Surface Sampling of Cleanrooms Bracketing (Suite 115) |  |  |  |  |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms |  | Weekly | 15Sep26 | SMO | Week on Testing Date |  |  | No growth | N/A | N/A |  | None |