Phase I

| # Section A – Test Details |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| # Initiator: | # Simin Mohammad (Written by: Maryam Naeem) | # Test Date: |  | # 07-May-2026 |  |  |  | # Date Initiated: |  | 18-May-2026 |  |
| # Date of Incident: | 18-May-2026 | # Test Name: |  | Environmental Monitoring |  |  |  | # Sample ID: |  | # ETX-260518-0254 |  |
| # Name of Analyst who Performed the Test: | # Simin Mohammad (Weekly Active Air Sampling Analyst) Samera A. Salim and Simin Mohammad (Weekly Active Air Sampling Plate Reader) | # Sample / Active Name |  | # EM SMO 116A Air 07MAY2026 |  |  |  | # Lot #: |  | # EM SMO 116A Air 07MAY2026 |  |
|  |  | # Dosage Form |  | # Plate |  |  |  |  |  |  |  |
| # Description of # Incident: | # The CFU count for the environmental monitoring plate exceeded the action level. |  |  |  | # Client Care notified of OOS # Sterility and Endotoxin test related OOS; Client must be notified within 24 hours of OOS |  |  |  |  |  |  |
|  | SOP / Test Method #: MICRO-SOP-2 | Effective Date: 05 Aug 2025 |  |  |  |  | SOP / Test Method Rev: 15 |  | # Limits/Specification: # <1 event |  |  |
| # Manager notified | # Manager/designee name: # Kathan Parikh | Date: 18-May-2026 |  |  |  |  | # Quality Assurance Department Notified: |  | # Yes No |  |  |
| # Section B – Investigation and Root Cause Analysis |  |  |  |  |  |  |  |  |  |  |  |
| # Personnel Discrepancy |  |  |  |  |  | # Comments |  |  |  |  |  |
| # Analyst interviewed? |  | # Yes No N/A |  |  |  | # Yes, analysts Simin Mohammad and Samera A. Salim and Simin Mohammad were comprehensively interviewed. |  |  |  |  |  |
| # Correct samples analyzed? |  | # Yes No N/A |  |  |  | # Yes, as per MICRO-SOP-2 |  |  |  |  |  |
| # Correct SOP / Test Method used/followed? |  | # Yes No N/A |  |  |  | Yes, as per MICRO-SOP-2 |  |  |  |  |  |
| # Correct sample collection/technique? |  | # Yes No N/A |  |  |  | Yes, as per MICRO-SOP-2 |  |  |  |  |  |
| # Are the Test Result Records Attached & Calculations Verified and Correct? |  | # Yes No N/A |  |  |  | # Yes, Information is available on Eagletrax under ETX-260518-0254 |  |  |  |  |  |
| # Is the analyst qualified to perform the test? |  | # Yes No N/A |  |  |  | Yes, the analysts are trained and qualified by quality to perform the test |  |  |  |  |  |
| # Is formula worksheet provided and verified? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  |  |
| # Does the processing method/ method set conform to the required specification? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  |  |
| # Was the sample stored appropriately? |  | # Yes No N/A |  |  |  | Yes, as per MICRO-SOP-2 |  |  |  |  |  |
| # Material Discrepancy |  |  |  |  |  | # Lot Number |  |  |  |  | # Expiry Date |
| # Is the integrity of the sample(s) questionable? |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |
| # Are the other test results from this run under investigation? |  | # Yes No N/A (See Phase I summary) |  |  |  |  |  |  |  |  |  |
| # Appropriate /correct glassware used? |  | # Yes No N/A |  |  |  |  |  |  |  |  |  |
| # Are reagents stored in proper condition and not expired? |  | # Yes No N/A |  |  |  | TSA Plate: 1011543730 |  |  |  |  | TSA Plate: 29SEP2026 |
| # Standard Used |  |  |  |  |  |  |  |  |  |  |  |
| # Solvent Used |  | # Not Applicable |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Is the standard the source of the root cause? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Has the analyst been instructed to keep the original test solution for possible retesting? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Equipment Discrepancy |  |  |  |  |  | # Equipment ID |  |  |  |  | # Calibration Due Date |
| # Was the instrument set up correctly? |  | # Yes No N/A |  |  |  | Please see below |  |  |  |  | Please see below |
| # Is the clean room facility certified? |  | Yes No N/A |  |  |  | CR116 (E001738) |  |  |  |  | December 2026 |
| # System Performance Check (if applicable)? |  | # Yes No N/A |  |  |  | N/A |  |  |  |  | N/A |
| # System Suitability RSD < 2% (PDA)? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # System Suitability RSD < 10% (QDA,CAD)? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Do the environmental monitoring results conform to required standards? |  | Yes No N/A |  |  |  | # See Phase I Summary |  |  |  |  | # See Phase I Summary |
| # Do the controls (i.e., positive, negative, etc.) conform to the required standards? |  | Yes No N/A |  |  |  | # See Phase I Summary |  |  |  |  | # See Phase I Summary |
| # Is equipment/instrument functional and calibrated/qualified/certified/PM? |  | # Yes No N/A |  |  |  | # Incubator E001034 # (Sensor E001501) # Incubator E001031 (Sensor E001505) |  |  |  |  | # Aug26 # Feb26 # Aug26 # Feb26 |
| # Are chromatogram/spectra/raw data correct? |  | # Yes No N/A |  |  |  | # Not Applicable |  |  |  |  | # Not Applicable |
| # Section C – Phase I Summary |  |  |  |  |  |  |  |  |  |  |  |
| # Has root cause been identified as a possible laboratory error? # Note: Root cause must be conclusive, verifiable, and well-documented. # Yes # If retesting is required, complete Phase II of the OOS (initiate Form 3.100.019.F02). # No, the OOS result is valid. |  |  |  |  |  |  |  |  |  |  |  |
| Phase I Summary: # Environmental Monitoring Summary: Weekly surface and active air sampling for CR116 for the previous week and following week of testing showed no microbial growth. However, the Weekly Active Air Sampling for CR116 for the week of testing, performed on 07 May 2026 by analyst SMO, exhibited 748 CFUs on EM SMO 116A Air 07MAY2026, which were identified as Kocuria palustris (Gram (+) cocci). Routine monitoring on the date of testing showed no growth across other monitored locations. During the interview with the analyst, they indicated that no obvious abnormalities or deviations in the testing procedure were observed. All materials were disinfected prior to testing. Moreover, the cleanroom suites in CR116 were thoroughly cleaned and prepared before initiating testing as per MICRO-SOP-2 and MICRO-SOP-9. Monthly cleaning and disinfection of the cleanroom suite, including the ISO 8 anteroom (116), ISO 7 buffer room (116A), ISO 7 cleanroom (116B), and the ISO 5 biosafety cabinets located within Room 116B, were performed on 30 Apr 2026 by Analyst - Rey Estrada - in accordance with MICRO-SOP-9, Cleaning and Disinfecting Procedure for Microbiology. All H2O2 indicators passed, confirming the successful completion and effectiveness of the monthly cleaning and disinfection activities within Rooms 116, 116A, and 116B. Additionally, routine cleaning and disinfection were performed before and after the testing activity in accordance with MICRO-SOP-9. It is important to note that no samples processed within that week in CR116 failed SCAN RDI testing that week. Based on the available evidence, the recovery of 748 CFU of Kocuria palustris (Gram (+) cocci) from EM SMO 116A Air 07MAY2026 in 116A on the date of testing (07 May 2026) appears to be an isolated event. This assessment is supported by the absence of microbial recovery from surrounding monitored areas and negative routine monitoring results throughout the bracketing period. The negative settling, personnel, and bracketing surface monitoring results demonstrate that the critical environment remained in a state of control. The available data do not support migration, persistence, or recurrence of contamination within CR116. Collectively, the evidence supports that the recovery was an isolated, transient, and non-recurring event, while established cleaning, disinfection, and aseptic controls remained effective. Accordingly, no systemic environmental control deficiencies were identified, and no additional corrective or preventive actions are warranted at this time beyond continued routine environmental monitoring and adherence to approved cleaning, disinfection, and aseptic procedures. # Accordingly, no systemic environmental control deficiencies were identified, and no additional corrective or preventive actions are warranted at this time beyond continued routine environmental monitoring and adherence to approved cleaning, disinfection, and aseptic procedures. |  |  |  |  |  |  |  |  |  |  |  |
| # Reviewer Comments: |  |  |  |  |  |  |  |  |  |  |  |
| # Acknowledgement and Closure |  |  |  |  |  |  |  |  |  |  |  |
| # Can investigation be closed? # Yes # Original result is valid. # No, initiate Form 3.100.019.F02. |  |  |  |  |  |  |  |  |  |  |  |
| # Prepared by (Print): |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |
| # Lab Manager/designee (Print): |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |
| # Approved by QA Representative (Print): # Closed: YES NO |  |  | # Signature/Date: |  |  |  |  |  |  |  |  |

# Table 1: Information for ETX-260518-0254 under investigation

# Table 2: Environmental Monitoring from Testing Performed on 07-May-2026

# Table 3: In Trend of Past OOS Results for Eagle Analytical Internal EM for analyte "EM SMO 116A Air 07MAY2026"

Table 1: Read Dates & Incubation Observation

Table 2: Environmental Monitoring for Analyst & Cleanroom

| ETX Submission ID | Set-up Analyst & Date | Sampling Site & Location | Plate Reading Analyst (≥ 48H) | CFUs Observed after 48 Hour Incubation at 30–35°C (E001031) | Plate Reading Analyst (NLT 5 days) | CFUs Observed after NLT 5-day Incubation at 20–25°C (E001034) | Microbial Identification |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ETX-260518-0254 | SMO 07-May-2026 | Weekly Active Air Sampling Plate (Cleanroom Suite 116A) | MC | No microbial growth was observed | SMO | 748 CFU on EM SMO 116A Air 07MAY2026 | Kocuria palustris (Gram (+) cocci) |

| Environmental Monitoring (EM) Sampling Site | Frequency | Date (DDMMMYYYY) | Analyst (Initials) | Day /Week(s) | Observation | Environmental Monitoring Plate ETX ID | Microbial ID | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Personnel EM Bracketing |  |  |  |  |  |  |  |  |
| Personal (Left Touch and Right Touch) | Daily | 06 May 2026 | SMO | Date Before Testing | No growth | N/A | N/A | None |
| Personal (Left Touch and Right Touch) | Daily | 07-May-2026 | SMO | Date of Testing | No growth | N/A | N/A | None |
| Personal (Left Touch and Right Touch) | Daily | 08 May 2026 | SMO | Date After Testing | No growth | N/A | N/A | None |
| Biological Safety Cabinet EM Bracketing Biological Safety Cabinet (BSC) |  |  |  |  |  |  |  |  |
| Surface Sampling of ISO 5 (4 locations) | Daily | 06 May 2026 | SMO | Date Before Testing | No growth | N/A | N/A | None |
| Surface Sampling of ISO 5 (4 locations) | Daily | 07-May-2026 | SMO | Date of Testing | No growth | N/A | N/A | None |
| Surface Sampling of ISO 5 (4 locations) | Daily | 08 May 2026 | SMO | Date After Testing | No growth | N/A | N/A | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 06 May 2026 | SMO | Date Before Testing | No growth | N/A | N/A | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 07-May-2026 | SMO | Date of Testing | No growth | N/A | N/A | None |
| Settling Sampling of ISO 5 (2 locations) | Daily | 08 May 2026 | SMO | Date After Testing | No growth | N/A | N/A | None |
| Weekly Active Air Sampling Bracketing |  |  |  |  |  |  |  |  |
| Active Air Sampling of Cleanrooms | Weekly | 06 May 2026 | Rey Estrada | Week (On or After Testing Date) | No growth | N/A | N/A | None |
| Surface Sampling of Anteroom and Cleanroom Bracketing |  |  |  |  |  |  |  |  |
| Surface Sampling of Cleanrooms | Weekly | 06 May 2026 | Rey Estrada | Week (On or After Testing Date) | No growth | N/A | N/A | None |